## raw
Test Label | Date | sample | component | rep | Mg Concentration | Al Concentration | Si Concentration | P Concentration | S Concentration | K Concentration | Ca Concentration | Ti Concentration | V Concentration | Cr Concentration | Mn Concentration | Fe Concentration | Co Concentration | Ni Concentration | Cu Concentration | Zn Concentration | As Concentration | Se Concentration | Rb Concentration | Sr Concentration | Y Concentration | Zr Concentration | Nb Concentration | Mo Concentration | Ag Concentration | Cd Concentration | Sn Concentration | Sb Concentration | Ba Concentration | W Concentration | Au Concentration | Hg Concentration | Pb Concentration | Bi Concentration | Th Concentration | U Concentration | LE Concentration
1 [Col] | 2023-06-06 | NIST 2711a | std | rep 1 | 15406 | 68395 | 261002 | 181 | 168 | 25265 | 23038 | 2921 | 78 | 48 | 656 | 29434 | <LOD | 27 | 710 | 378 | 150 | 7 | 119 | 268 | 31 | 281 | 18 | 23 | 40 | 105 | 49 | 125 | 1076 | <LOD | <LOD | 10 | 1378 | 24 | <LOD | <LOD | 728608
2 [Col] | 2023-06-06 | NIST 2711a | std | rep 2 | 15196 | 68055 | 257114 | 177 | 246 | 25100 | 22126 | 3961 | <LOD | 39 | 742 | 29897 | <LOD | 33 | 712 | 387 | 135 | <LOD | 123 | 263 | 29 | 263 | 16 | 7 | 16 | 76 | 42 | 57 | 1067 | <LOD | 6 | 20 | 1359 | <LOD | 18 | <LOD | 731904
3 [Col] | 2023-06-06 | NIST 2711a | std | rep 3 | 16362 | 68292 | 261493 | 85 | 245 | 24998 | 22349 | 3169 | 142 | 63 | 657 | 29356 | 148 | 42 | 699 | 383 | 74 | 4 | 124 | 262 | 29 | 306 | 24 | <LOD | 11 | 78 | 82 | 107 | 1201 | <LOD | 5 | 19 | 1407 | 33 | <LOD | <LOD | 727237
4 [Col] | 2023-06-06 | NIST 2711a | std | rep 4 | 14584 | 68798 | 258313 | 243 | 221 | 25170 | 22990 | 3126 | 186 | 95 | 684 | 28880 | <LOD | 36 | 727 | 341 | 102 | 2 | 121 | 259 | 39 | 310 | 24 | 13 | 14 | 74 | 53 | 132 | 655 | 44 | <LOD | 22 | 1347 | <LOD | 22 | <LOD | 732344
5 [Col] | 2023-06-06 | NIST 2711a | std | rep 5 | 16346 | 68521 | 262451 | 176 | 257 | 25769 | 22725 | 2834 | <LOD | 97 | 651 | 29052 | 122 | 21 | 718 | 398 | 145 | 3 | 120 | 257 | 32 | 294 | 13 | 4 | 39 | 89 | 76 | 140 | 1241 | 17 | <LOD | 14 | 1353 | <LOD | 11 | 4 | 726108
6 [Col] | 2023-06-06 | NIST 679 | std | rep 1 | 11618 | 90098 | 218211 | <LOD | <LOD | 23642 | 1554 | 5272 | 213 | 104 | 1353 | 91626 | 218 | 61 | 908 | 108 | 24 | 10 | 178 | 51 | 47 | 170 | 20 | 10 | 43 | 76 | 101 | 126 | 534 | 32 | <LOD | <LOD | 12 | <LOD | <LOD | 16 | 706587
7 [Col] | 2023-06-06 | NIST 679 | std | rep 2 | 10844 | 90719 | 217962 | <LOD | <LOD | 23408 | 1303 | 5395 | 240 | 129 | 1561 | 89718 | 470 | 41 | 911 | 101 | 19 | 6 | 186 | 45 | 45 | 157 | 20 | 13 | 19 | 50 | <LOD | 49 | <LOD | <LOD | 9 | 8 | 20 | <LOD | 30 | <LOD | 709596
8 [Col] | 2023-06-06 | NIST 679 | std | rep 3 | 11975 | 90447 | 217339 | <LOD | <LOD | 23538 | 1495 | 5457 | 158 | 78 | 1457 | 88786 | 385 | 56 | 952 | 132 | 17 | 6 | 192 | 54 | 36 | 154 | 19 | 14 | 19 | 27 | 26 | 71 | 504 | 17 | <LOD | 6 | 17 | <LOD | 20 | <LOD | 709443
9 [Col] | 2023-06-06 | NIST 679 | std | rep 4 | 11771 | 87623 | 213484 | <LOD | <LOD | 23108 | 1492 | 4604 | 139 | 31 | 1521 | 90073 | 195 | 81 | 999 | 108 | 28 | 6 | 196 | 44 | 45 | 166 | 20 | 9 | 31 | <LOD | 22 | <LOD | 606 | <LOD | <LOD | 7 | 4 | <LOD | 25 | 11 | 715442
10 [Col] | 2023-06-06 | NIST 679 | std | rep 5 | 12112 | 90489 | 217735 | <LOD | <LOD | 23431 | 1349 | 5363 | 138 | 107 | 1393 | 90512 | 268 | 55 | 900 | 109 | 21 | 2 | 194 | 53 | 43 | 154 | 17 | 12 | 35 | 30 | 59 | 90 | 1030 | 57 | <LOD | 8 | 16 | <LOD | 34 | <LOD | 707007
11 [Col] | 2023-06-06 | SARM 69 | std | rep 1 | 13709 | 76526 | 257506 | 327 | <LOD | 17075 | 14308 | 4872 | 89 | 202 | 922 | 50482 | 183 | 77 | 720 | 62 | 12 | 3 | 72 | 110 | 29 | 286 | 7 | <LOD | 18 | 15 | 35 | 72 | 1258 | 15 | 5 | 9 | 10 | <LOD | 11 | <LOD | 716059
12 [Col] | 2023-06-06 | SARM 69 | std | rep 2 | 13622 | 75672 | 256237 | 275 | <LOD | 17090 | 14055 | 4582 | 157 | 230 | 950 | 50338 | 184 | 82 | 717 | 70 | 6 | 2 | 73 | 106 | 30 | 258 | <LOD | <LOD | 32 | 26 | 42 | 74 | 808 | <LOD | 5 | <LOD | 15 | 23 | <LOD | 6 | 718977
13 [Col] | 2023-06-06 | SARM 69 | std | rep 3 | 13780 | 78657 | 263033 | 378 | <LOD | 17728 | 14219 | 4894 | 73 | 230 | 892 | 51263 | 81 | 60 | 695 | 72 | 5 | 2 | 76 | 108 | 35 | 295 | 8 | 8 | 14 | 18 | 39 | 102 | 1016 | <LOD | 8 | 13 | 22 | <LOD | 30 | <LOD | 708343
14 [Col] | 2023-06-06 | SARM 69 | std | rep 4 | 14013 | 77359 | 259477 | 380 | <LOD | 17389 | 13905 | 4856 | 278 | 235 | 900 | 51579 | 66 | 67 | 680 | 58 | 6 | 3 | 76 | 100 | 29 | 338 | 16 | 11 | 18 | <LOD | 29 | 68 | 1351 | 16 | <LOD | 5 | 17 | <LOD | 11 | <LOD | 712015
15 [Col] | 2023-06-06 | SARM 69 | std | rep 5 | 14286 | 77760 | 260453 | 313 | <LOD | 17061 | 13797 | 4280 | 145 | 197 | 979 | 50676 | 258 | 92 | 656 | 70 | 6 | 3 | 69 | 97 | 30 | 271 | 14 | 16 | 37 | 35 | 31 | 72 | 2076 | <LOD | 6 | 6 | 18 | <LOD | 12 | 7 | 711496
1 [Col] | 2023-06-07 | NIST 2711a | std | rep 1 | 14252 | 68263 | 262212 | 119 | 130 | 25592 | 23173 | 3722 | <LOD | 59 | 781 | 29821 | 188 | 38 | 698 | 399 | 142 | 3 | 135 | 266 | 31 | 312 | 25 | <LOD | 28 | 81 | 30 | 122 | 889 | <LOD | 5 | 22 | 1438 | <LOD | 14 | <LOD | 727393
2 [Col] | 2023-06-07 | NIST 2711a | std | rep 2 | 15171 | 68069 | 258473 | 196 | 152 | 24445 | 21616 | 3077 | 144 | 101 | 708 | 28677 | <LOD | 33 | 692 | 360 | 146 | 4 | 118 | 243 | 32 | 305 | 24 | 9 | <LOD | 68 | 44 | 44 | 1507 | 21 | 6 | <LOD | 1259 | 18 | <LOD | 10 | 733148
3 [Col] | 2023-06-07 | NIST 2711a | std | rep 3 | 15056 | 69390 | 258678 | 223 | 170 | 25575 | 22951 | 3202 | 103 | 88 | 675 | 28611 | 63 | 31 | 734 | 378 | 115 | <LOD | 119 | 248 | 34 | 324 | 16 | 9 | 19 | 77 | 57 | 68 | 1134 | <LOD | 9 | 14 | 1350 | <LOD | 23 | <LOD | 730645
4 [Col] | 2023-06-07 | NIST 2711a | std | rep 4 | 14898 | 69475 | 262233 | 81 | 123 | 25742 | 22588 | 3045 | 117 | 122 | 687 | 29967 | 66 | 28 | 716 | 411 | 133 | 4 | 122 | 271 | 24 | 296 | 24 | 7 | 23 | 104 | 33 | 74 | 662 | <LOD | 9 | 10 | 1382 | <LOD | 12 | <LOD | 726953
5 [Col] | 2023-06-07 | NIST 2711a | std | rep 5 | 15312 | 68777 | 258201 | 202 | 202 | 25197 | 22394 | 3651 | 150 | 80 | 765 | 28090 | 140 | 36 | 707 | 402 | 128 | 3 | 119 | 260 | 34 | 349 | 16 | 17 | 20 | 56 | 40 | 84 | 1111 | <LOD | 11 | 9 | 1313 | <LOD | 13 | <LOD | 731690
6 [Col] | 2023-06-07 | NIST 679 | std | rep 1 | 10864 | 93377 | 222631 | <LOD | <LOD | 24106 | 1434 | 5696 | 218 | 121 | 1535 | 91622 | 236 | 45 | 844 | 117 | 20 | 10 | 187 | 51 | 42 | 163 | 18 | 10 | 32 | 43 | 58 | 52 | 698 | 18 | <LOD | <LOD | 14 | <LOD | 16 | 9 | 700035
7 [Col] | 2023-06-07 | NIST 679 | std | rep 2 | 12601 | 90330 | 217378 | <LOD | <LOD | 23658 | 1471 | 5320 | 118 | 87 | 1640 | 89016 | 241 | 53 | 925 | 135 | 31 | 8 | 185 | 49 | 33 | 159 | 28 | 26 | <LOD | 16 | 57 | 64 | 630 | 21 | 4 | 10 | 9 | <LOD | 15 | <LOD | 708503
8 [Col] | 2023-06-07 | NIST 679 | std | rep 3 | 10807 | 90790 | 218912 | <LOD | <LOD | 23302 | 1411 | 4882 | 177 | 92 | 1603 | 93207 | <LOD | 68 | 843 | 124 | 16 | 6 | 194 | 61 | 37 | 169 | 24 | 16 | <LOD | 21 | 33 | 73 | 1200 | <LOD | 5 | 12 | 15 | <LOD | 29 | <LOD | 705051
9 [Col] | 2023-06-07 | NIST 679 | std | rep 4 | 11346 | 87036 | 210294 | <LOD | <LOD | 22146 | 1382 | 5142 | 116 | 167 | 1368 | 82667 | 297 | 40 | 762 | 117 | 23 | 4 | 167 | 44 | 43 | 151 | 14 | <LOD | 50 | 68 | 42 | 40 | 617 | <LOD | 6 | <LOD | 9 | 41 | <LOD | <LOD | 726956
10 [Col] | 2023-06-07 | NIST 679 | std | rep 5 | 11895 | 90440 | 219276 | <LOD | <LOD | 23551 | 1400 | 4952 | 146 | 145 | 1733 | 93563 | 148 | 66 | 894 | 135 | 26 | 2 | 197 | 61 | 39 | 164 | 22 | 12 | <LOD | <LOD | 34 | 79 | 1423 | 30 | <LOD | 8 | 14 | <LOD | 12 | <LOD | 702574
11 [Col] | 2023-06-07 | SARM 69 | std | rep 1 | 13621 | 77075 | 261201 | 392 | <LOD | 17468 | 13843 | 4173 | 298 | 203 | 909 | 50482 | 70 | 48 | 653 | 51 | 11 | 3 | 73 | 96 | 28 | 347 | 16 | 6 | 16 | 33 | 58 | 60 | 470 | 72 | <LOD | 7 | 12 | 44 | <LOD | 8 | 713708
12 [Col] | 2023-06-07 | SARM 69 | std | rep 2 | 13424 | 76567 | 262417 | 275 | <LOD | 17404 | 13992 | 4667 | 95 | 227 | 978 | 51539 | 230 | 82 | 639 | 63 | 10 | 4 | 79 | 102 | 25 | 247 | 9 | <LOD | <LOD | 35 | 63 | 46 | 799 | 29 | <LOD | 6 | 14 | <LOD | 8 | <LOD | 711467
13 [Col] | 2023-06-07 | SARM 69 | std | rep 3 | 13145 | 78056 | 261324 | 323 | <LOD | 17660 | 14190 | 4728 | 132 | 250 | 1004 | 51741 | 182 | 49 | 667 | 59 | 8 | 7 | 79 | 94 | 30 | 272 | 9 | 8 | 27 | 14 | 60 | 96 | 1262 | <LOD | 4 | <LOD | 17 | <LOD | <LOD | <LOD | 710513
14 [Col] | 2023-06-07 | SARM 69 | std | rep 4 | 13194 | 78612 | 266546 | 328 | <LOD | 17039 | 14030 | 4593 | 136 | 275 | 952 | 50942 | 312 | 74 | 615 | 71 | 11 | 5 | 73 | 103 | 28 | 282 | 10 | <LOD | 23 | 47 | 70 | 34 | 605 | 30 | <LOD | <LOD | 16 | <LOD | 16 | <LOD | 707152
15 [Col] | 2023-06-07 | SARM 69 | std | rep 5 | 11923 | 78182 | 265931 | 366 | <LOD | 17292 | 14315 | 4276 | 173 | 262 | 985 | 51250 | 162 | 46 | 663 | 71 | 9 | 7 | 75 | 110 | 27 | 293 | 8 | 7 | 36 | 58 | 55 | 51 | 1531 | <LOD | <LOD | 6 | 15 | <LOD | 14 | 7 | 708323
2 [Col] | 2023-06-08 | NIST 2711a | std | rep 1 | 14108 | 68765 | 261965 | 147 | 216 | 25611 | 22608 | 3145 | 85 | 38 | 655 | 28830 | 103 | 47 | 555 | 370 | 117 | 4 | 125 | 271 | 33 | 319 | 20 | 9 | <LOD | 76 | 74 | 98 | 1772 | 53 | <LOD | 6 | 1389 | <LOD | 9 | 5 | 728724
3 [Col] | 2023-06-08 | NIST 2711a | std | rep 2 | 14786 | 67043 | 258900 | 193 | 199 | 25394 | 22975 | 3231 | 135 | 42 | 749 | 28990 | 204 | 26 | 587 | 414 | 144 | <LOD | 124 | 249 | 31 | 314 | 20 | 5 | 23 | 59 | 37 | 37 | 1229 | 23 | <LOD | <LOD | 1416 | <LOD | 29 | <LOD | 732041
4 [Col] | 2023-06-08 | NIST 2711a | std | rep 3 | 15057 | 67779 | 261198 | 200 | 149 | 24613 | 22718 | 3586 | 119 | 143 | 654 | 29332 | 82 | 24 | 543 | 406 | 123 | 3 | 115 | 259 | 32 | 291 | 21 | 16 | 35 | 88 | 43 | 63 | 1029 | 28 | <LOD | 8 | 1310 | <LOD | 25 | <LOD | 729430
5 [Col] | 2023-06-08 | NIST 2711a | std | rep 4 | 15533 | 67327 | 255281 | 205 | 184 | 24654 | 22468 | 3363 | 75 | 126 | 561 | 28655 | <LOD | 30 | 568 | 380 | 125 | 3 | 115 | 258 | 32 | 287 | 21 | 9 | 22 | 95 | 70 | 85 | 1184 | 48 | <LOD | 8 | 1296 | 26 | <LOD | <LOD | 735741
6 [Col] | 2023-06-08 | NIST 2711a | std | rep 5 | 15071 | 68115 | 257607 | 184 | 161 | 24845 | 22090 | 3083 | 202 | 153 | 643 | 28079 | 137 | 43 | 546 | 404 | 149 | 3 | 116 | 245 | 33 | 315 | 17 | 10 | <LOD | 71 | 48 | 91 | 629 | 17 | <LOD | 8 | 1315 | <LOD | 17 | <LOD | 734761
7 [Col] | 2023-06-08 | NIST 679 | std | rep 1 | 9769 | 92172 | 219291 | <LOD | <LOD | 23535 | 1384 | 5409 | 220 | <LOD | 1603 | 89537 | 165 | 52 | 593 | 100 | 22 | 2 | 179 | 41 | 31 | 156 | 16 | 12 | 32 | 58 | 60 | 88 | 634 | 58 | <LOD | <LOD | 12 | <LOD | 18 | <LOD | 708506
8 [Col] | 2023-06-08 | NIST 679 | std | rep 2 | 11661 | 94095 | 222594 | <LOD | <LOD | 24338 | 1427 | 5180 | 210 | 94 | 1676 | 91289 | 265 | 64 | 624 | 126 | 24 | 5 | 194 | 49 | 33 | 159 | 23 | 11 | 28 | 48 | 62 | 94 | <LOD | 62 | <LOD | <LOD | 17 | <LOD | 41 | <LOD | 699973
9 [Col] | 2023-06-08 | NIST 679 | std | rep 3 | 12091 | 94632 | 224001 | <LOD | <LOD | 24182 | 1487 | 5177 | 143 | 109 | 1733 | 92700 | 240 | 63 | 640 | 121 | 24 | 6 | 192 | 48 | 32 | 156 | 24 | <LOD | <LOD | 15 | 53 | 71 | <LOD | <LOD | 6 | <LOD | 4 | <LOD | 10 | 5 | 696609
10 [Col] | 2023-06-08 | NIST 679 | std | rep 4 | 11492 | 92810 | 222277 | <LOD | <LOD | 23724 | 1308 | 5678 | 127 | 102 | 1565 | 91513 | 326 | 71 | 629 | 130 | 18 | 5 | 182 | 51 | 44 | 176 | 19 | 15 | 47 | 41 | 60 | 79 | 854 | 40 | <LOD | <LOD | 18 | <LOD | 24 | <LOD | 700429
11 [Col] | 2023-06-08 | NIST 679 | std | rep 5 | 12289 | 91531 | 220863 | <LOD | <LOD | 23832 | 1357 | 5010 | 160 | 118 | 1557 | 91485 | 183 | 50 | 598 | 118 | 23 | 7 | 198 | 53 | 39 | 163 | 11 | 14 | 40 | 68 | 60 | 164 | 605 | <LOD | <LOD | 7 | 12 | <LOD | 36 | <LOD | 702800
12 [Col] | 2023-06-08 | SARM 69 | std | rep 1 | 14791 | 78606 | 265484 | 425 | <LOD | 17717 | 13876 | 4504 | 314 | 199 | 1070 | 51427 | 235 | 51 | 518 | 58 | 5 | 2 | 77 | 106 | 33 | 321 | 12 | 7 | 18 | 25 | 55 | 80 | 878 | 39 | <LOD | <LOD | 15 | <LOD | 9 | 3 | 705151
13 [Col] | 2023-06-08 | SARM 69 | std | rep 2 | 14941 | 76553 | 259836 | 330 | <LOD | 17282 | 13434 | 4876 | 212 | 245 | 826 | 51466 | 179 | 59 | 536 | 54 | 17 | 2 | 77 | 108 | 24 | 241 | 8 | 6 | 29 | 50 | 46 | <LOD | 857 | 25 | <LOD | <LOD | 7 | <LOD | 21 | 5 | 712479
14 [Col] | 2023-06-08 | SARM 69 | std | rep 3 | 13904 | 79413 | 267378 | 365 | <LOD | 17716 | 14706 | 4364 | 196 | 207 | 922 | 51355 | 242 | 60 | 534 | 65 | 6 | 7 | 77 | 95 | 32 | 314 | 12 | 9 | 14 | 35 | 66 | 64 | 1029 | <LOD | 4 | 7 | 14 | <LOD | 15 | <LOD | 703670
15 [Col] | 2023-06-08 | SARM 69 | std | rep 4 | 12686 | 76793 | 264770 | 344 | <LOD | 17444 | 14059 | 4701 | 309 | 176 | 1181 | 52423 | <LOD | 53 | 536 | 71 | 7 | 2 | 82 | 113 | 28 | 294 | 7 | 15 | <LOD | 73 | 58 | 72 | 806 | <LOD | 8 | 7 | 15 | 16 | <LOD | <LOD | 708787
16 [Col] | 2023-06-08 | SARM 69 | std | rep 5 | 13693 | 78361 | 263921 | 369 | <LOD | 17205 | 13842 | 4304 | 198 | 147 | 889 | 50413 | <LOD | 46 | 492 | 78 | 4 | 9 | 78 | 102 | 27 | 254 | 8 | 6 | 32 | 46 | 66 | 47 | 622 | <LOD | <LOD | 10 | 25 | <LOD | <LOD | <LOD | 710613
1 [Col] | 2023-06-09 | NIST 2711a | std | rep 1 | 13434 | 67975 | 256138 | 203 | 267 | 24449 | 22510 | 3056 | 92 | 74 | 645 | 27423 | <LOD | 24 | 595 | 384 | 145 | <LOD | 116 | 244 | 29 | 277 | 21 | 9 | 21 | 87 | 51 | 83 | 1697 | <LOD | 7 | 13 | 1296 | <LOD | 13 | 8 | 737954
2 [Col] | 2023-06-09 | NIST 2711a | std | rep 2 | 15238 | 68918 | 259395 | 140 | 114 | 25333 | 22224 | 3895 | 113 | 33 | 692 | 28961 | 65 | 30 | 555 | 387 | 111 | 6 | 123 | 277 | 43 | 290 | 14 | 6 | 19 | 82 | 50 | 48 | 975 | <LOD | 6 | 10 | 1353 | <LOD | 8 | <LOD | 730173
3 [Col] | 2023-06-09 | NIST 2711a | std | rep 3 | 14197 | 67748 | 257255 | 190 | 132 | 24678 | 22375 | 2726 | 75 | 74 | 589 | 28691 | 65 | 24 | 559 | 396 | 146 | <LOD | 114 | 254 | 30 | 317 | 15 | 6 | 17 | 82 | 61 | 115 | 1457 | 38 | <LOD | 6 | 1317 | <LOD | <LOD | <LOD | 735564
4 [Col] | 2023-06-09 | NIST 2711a | std | rep 4 | 15624 | 70506 | 261592 | 166 | 214 | 25141 | 23285 | 3491 | <LOD | 165 | 696 | 29608 | 58 | 37 | 645 | 406 | 140 | 5 | 128 | 271 | 31 | 339 | 20 | <LOD | 15 | 106 | 31 | 79 | 1518 | <LOD | 10 | 20 | 1352 | <LOD | <LOD | <LOD | 724786
5 [Col] | 2023-06-09 | NIST 2711a | std | rep 5 | 16523 | 71054 | 262067 | 217 | 201 | 25713 | 22431 | 3233 | 66 | 124 | 715 | 29267 | 92 | 27 | 560 | 394 | 124 | 2 | 118 | 254 | 31 | 308 | 20 | 4 | 40 | 124 | 86 | 77 | 1404 | <LOD | <LOD | 12 | 1384 | <LOD | <LOD | <LOD | 723754
6 [Col] | 2023-06-09 | NIST 679 | std | rep 1 | 9769 | 91309 | 219204 | <LOD | <LOD | 23521 | 1449 | 4883 | 158 | 90 | 1533 | 88714 | <LOD | 66 | 593 | 109 | 19 | 7 | 191 | 50 | 35 | 156 | 27 | 21 | 29 | 20 | 27 | 83 | 796 | <LOD | 9 | <LOD | 20 | 22 | <LOD | <LOD | 710706
7 [Col] | 2023-06-09 | NIST 679 | std | rep 2 | 10882 | 91716 | 220485 | <LOD | <LOD | 24023 | 1501 | 5414 | 181 | 137 | 1648 | 92110 | <LOD | 48 | 652 | 133 | 37 | 7 | 191 | 59 | 43 | 163 | 24 | 23 | 15 | 43 | 62 | 63 | <LOD | 42 | <LOD | <LOD | 8 | <LOD | 25 | <LOD | 704058
8 [Col] | 2023-06-09 | NIST 679 | std | rep 3 | 12055 | 94045 | 220807 | <LOD | <LOD | 23708 | 1427 | 5335 | 199 | 82 | 1653 | 91603 | 166 | 53 | 597 | 137 | 32 | 7 | 176 | 47 | 39 | 173 | 27 | <LOD | <LOD | 42 | 32 | 27 | 1230 | <LOD | <LOD | 6 | 12 | <LOD | 12 | <LOD | 700255
9 [Col] | 2023-06-09 | NIST 679 | std | rep 4 | 11354 | 92525 | 223167 | <LOD | <LOD | 24288 | 1402 | 5109 | 125 | 131 | 1624 | 90911 | 261 | 51 | 621 | 116 | 14 | 3 | 204 | 56 | 38 | 164 | 25 | 11 | 55 | 60 | 71 | 95 | 708 | 33 | 4 | <LOD | 13 | <LOD | 23 | <LOD | 700937
10 [Col] | 2023-06-09 | NIST 679 | std | rep 5 | 12477 | 88158 | 212333 | <LOD | <LOD | 22390 | 1347 | 5578 | 188 | 89 | 1518 | 84854 | <LOD | 50 | 605 | 140 | 22 | 9 | 163 | 37 | 41 | 138 | 18 | 4 | <LOD | 31 | 87 | 106 | 671 | <LOD | <LOD | 6 | 11 | <LOD | <LOD | <LOD | 720337
11 [Col] | 2023-06-09 | SARM 69 | std | rep 1 | 12303 | 76517 | 261536 | 293 | <LOD | 17335 | 13750 | 4693 | <LOD | 174 | 921 | 51086 | 130 | 53 | 523 | 73 | 7 | 4 | 84 | 114 | 37 | 290 | 11 | 18 | <LOD | 24 | 60 | 81 | 1198 | 19 | <LOD | <LOD | 18 | <LOD | 15 | <LOD | 714169
12 [Col] | 2023-06-09 | SARM 69 | std | rep 2 | 13268 | 77329 | 259925 | 366 | <LOD | 17340 | 13937 | 5219 | <LOD | 252 | 875 | 50994 | 242 | 78 | 528 | 49 | 5 | 4 | 72 | 103 | 33 | 322 | 12 | 7 | 18 | 33 | 66 | 63 | 658 | 21 | <LOD | 8 | 21 | 21 | <LOD | <LOD | 713600
13 [Col] | 2023-06-09 | SARM 69 | std | rep 3 | 12999 | 78872 | 264625 | 303 | <LOD | 17928 | 14099 | 4750 | <LOD | 190 | 917 | 52254 | 173 | 72 | 494 | 68 | 6 | 4 | 75 | 104 | 35 | 314 | 10 | <LOD | 12 | <LOD | 35 | 51 | 1003 | <LOD | 8 | <LOD | 22 | <LOD | 18 | 8 | 707092
14 [Col] | 2023-06-09 | SARM 69 | std | rep 4 | 12894 | 77526 | 263657 | 319 | <LOD | 17470 | 14275 | 4881 | <LOD | 248 | 853 | 51818 | 99 | 51 | 538 | 81 | 8 | 3 | 78 | 105 | 35 | 324 | 6 | 7 | <LOD | 29 | 71 | 58 | <LOD | <LOD | 4 | <LOD | 15 | <LOD | 10 | <LOD | 710595
15 [Col] | 2023-06-09 | SARM 69 | std | rep 5 | 12982 | 78326 | 264064 | 295 | <LOD | 17553 | 13880 | 5213 | 123 | 198 | 850 | 51862 | 68 | 94 | 544 | 62 | 7 | <LOD | 77 | 105 | 31 | 334 | 9 | 23 | <LOD | <LOD | 36 | 53 | 877 | 30 | <LOD | 9 | 19 | <LOD | 12 | <LOD | 708375
1 [Col] | 2023-06-13 | NIST 2711a | std | rep 1 | 15464 | 72482 | 266609 | 276 | 189 | 26199 | 23712 | 3797 | 204 | 48 | 657 | 29081 | 91 | 37 | 443 | 403 | 119 | 3 | 119 | 269 | 35 | 342 | 20 | 5 | <LOD | 72 | 76 | 86 | 769 | <LOD | <LOD | 6 | 1362 | <LOD | 7 | 3 | 718921
2 [Col] | 2023-06-13 | NIST 2711a | std | rep 2 | 15788 | 70680 | 266770 | 121 | 114 | 25453 | 22690 | 3156 | 91 | 133 | 593 | 28839 | 76 | 32 | 378 | 368 | 138 | <LOD | 122 | 256 | 33 | 297 | 18 | 5 | 40 | 101 | 77 | 99 | 913 | <LOD | <LOD | 19 | 1286 | <LOD | 23 | 5 | 722130
3 [Col] | 2023-06-13 | NIST 2711a | std | rep 3 | 16657 | 70577 | 264937 | 187 | 210 | 25445 | 23440 | 3754 | 131 | <LOD | 639 | 28768 | <LOD | 23 | 407 | 396 | 127 | 2 | 121 | 249 | 28 | 279 | 18 | 5 | 35 | 103 | 44 | 58 | 1202 | <LOD | 10 | 18 | 1329 | <LOD | <LOD | 6 | 721553
4 [Col] | 2023-06-13 | NIST 2711a | std | rep 4 | 15610 | 69751 | 265080 | 127 | 112 | 25590 | 22007 | 2989 | <LOD | 100 | 695 | 28978 | 71 | 20 | 335 | 386 | 154 | 3 | 124 | 258 | 28 | 310 | 17 | 14 | <LOD | 59 | 43 | 115 | 1436 | 16 | <LOD | 15 | 1323 | 33 | 8 | <LOD | 724535
5 [Col] | 2023-06-13 | NIST 2711a | std | rep 5 | 15249 | 70785 | 265107 | 178 | 229 | 25311 | 23296 | 3336 | <LOD | 105 | 688 | 28114 | 111 | 24 | 351 | 384 | 124 | 5 | 117 | 239 | 27 | 273 | 23 | 13 | <LOD | 49 | 62 | 90 | 1241 | <LOD | <LOD | 10 | 1324 | 19 | <LOD | 5 | 724083
6 [Col] | 2023-06-13 | NIST 679 | std | rep 1 | 12623 | 91028 | 218799 | <LOD | <LOD | 23414 | 1478 | 5520 | 215 | 108 | 1742 | 90490 | <LOD | 53 | 415 | 107 | 21 | 4 | 182 | 51 | 42 | 178 | 19 | <LOD | 44 | 66 | 60 | <LOD | 1927 | <LOD | 5 | 8 | 13 | 40 | <LOD | <LOD | 704317
7 [Col] | 2023-06-13 | NIST 679 | std | rep 2 | 11756 | 92835 | 221562 | <LOD | <LOD | 23876 | 1416 | 5594 | 153 | 76 | 1631 | 90424 | 172 | 62 | 347 | 127 | 21 | 7 | 181 | 37 | 44 | 181 | 21 | <LOD | <LOD | 46 | 59 | 62 | 559 | <LOD | 8 | <LOD | 8 | <LOD | 26 | <LOD | 702581
8 [Col] | 2023-06-13 | NIST 679 | std | rep 3 | 9769 | 87924 | 215456 | <LOD | <LOD | 23255 | 1435 | 5409 | 142 | 136 | 1705 | 90164 | <LOD | 67 | 406 | 136 | 24 | 5 | 189 | 53 | 47 | 162 | 17 | 9 | <LOD | <LOD | 64 | 94 | 637 | 35 | <LOD | 7 | 13 | <LOD | 26 | 12 | 715017
9 [Col] | 2023-06-13 | NIST 679 | std | rep 4 | 9769 | 92004 | 222606 | <LOD | <LOD | 24295 | 1438 | 5548 | <LOD | 93 | 1578 | 93798 | 327 | 65 | 422 | 110 | 24 | 6 | 194 | 48 | 37 | 169 | 23 | 10 | 20 | 54 | 69 | 50 | 1238 | <LOD | <LOD | 11 | 15 | <LOD | 32 | 6 | 700214
10 [Col] | 2023-06-13 | NIST 679 | std | rep 5 | 11348 | 93862 | 222884 | <LOD | <LOD | 24279 | 1634 | 5119 | 195 | 84 | 1731 | 91793 | 155 | 57 | 402 | 120 | 28 | 8 | 184 | 50 | 42 | 167 | 21 | 15 | <LOD | 63 | 54 | 56 | 480 | <LOD | 8 | 5 | 4 | <LOD | 24 | <LOD | 699677
11 [Col] | 2023-06-13 | SARM 69 | std | rep 1 | 13795 | 76706 | 265469 | 331 | <LOD | 16968 | 14217 | 4503 | 163 | 224 | 1025 | 50699 | 96 | 42 | 321 | 66 | 16 | 2 | 72 | 96 | 28 | 292 | 10 | 12 | <LOD | <LOD | 59 | 66 | 867 | 25 | <LOD | 14 | 9 | <LOD | 13 | <LOD | 709473
12 [Col] | 2023-06-13 | SARM 69 | std | rep 2 | 13517 | 78899 | 264042 | 353 | <LOD | 17398 | 14058 | 4596 | 153 | 172 | 961 | 51363 | <LOD | 72 | 338 | 62 | 7 | 6 | 70 | 105 | 29 | 315 | 4 | 8 | 22 | 19 | 29 | 41 | 702 | 52 | <LOD | <LOD | 11 | <LOD | <LOD | 5 | 708791
14 [Col] | 2023-06-13 | SARM 69 | std | rep 3 | 13063 | 77670 | 266629 | 319 | <LOD | 17509 | 13998 | 3632 | 76 | 218 | 999 | 51203 | 177 | 48 | 319 | 63 | 7 | 7 | 75 | 98 | 38 | 279 | 9 | 12 | <LOD | <LOD | 73 | 41 | 1032 | <LOD | <LOD | <LOD | 17 | 23 | <LOD | <LOD | 708665
15 [Col] | 2023-06-13 | SARM 69 | std | rep 4 | 13155 | 77484 | 265664 | 368 | <LOD | 17017 | 14467 | 4139 | 153 | 252 | 1009 | 51309 | 151 | 66 | 322 | 61 | 9 | 5 | 71 | 104 | 30 | 323 | 18 | 5 | <LOD | <LOD | <LOD | 91 | 927 | <LOD | <LOD | 10 | 14 | <LOD | <LOD | 10 | 708875
16 [Col] | 2023-06-13 | SARM 69 | std | rep 5 | 13995 | 79465 | 266676 | 340 | <LOD | 17609 | 15093 | 5014 | 148 | 207 | 1008 | 52095 | 168 | 71 | 315 | 82 | 7 | 5 | 75 | 93 | 30 | 276 | 6 | 11 | <LOD | 34 | 48 | 127 | 597 | <LOD | <LOD | <LOD | 12 | <LOD | <LOD | 5 | 703267
1 [Col] | 2023-06-14 | NIST 2711a | std | rep 1 | 14514 | 69621 | 260878 | 152 | 209 | 25806 | 22636 | 3500 | 151 | 57 | 634 | 29837 | 112 | 26 | 368 | 396 | 105 | 5 | 122 | 260 | 37 | 269 | 15 | 4 | <LOD | 75 | 31 | 75 | 813 | <LOD | <LOD | 12 | 1358 | <LOD | <LOD | <LOD | 728369
2 [Col] | 2023-06-14 | NIST 2711a | std | rep 2 | 13762 | 69277 | 261147 | 167 | 178 | 25040 | 22754 | 2893 | <LOD | <LOD | 670 | 29289 | 192 | 32 | 377 | 400 | 161 | 5 | 125 | 266 | 39 | 292 | 21 | 15 | <LOD | 57 | 76 | 62 | 2350 | 50 | <LOD | <LOD | 1359 | 17 | <LOD | <LOD | 729215
3 [Col] | 2023-06-14 | NIST 2711a | std | rep 3 | 15094 | 70012 | 260175 | 158 | 98 | 25572 | 23570 | 3349 | <LOD | <LOD | 637 | 29014 | 102 | 10 | 389 | 395 | 138 | 4 | 117 | 244 | 31 | 271 | 12 | 8 | 11 | 73 | 26 | 68 | 1421 | 34 | <LOD | 7 | 1368 | <LOD | 18 | <LOD | 728223
4 [Col] | 2023-06-14 | NIST 2711a | std | rep 4 | 14423 | 68886 | 261413 | 124 | 116 | 24787 | 22654 | 3525 | 86 | 147 | 622 | 27924 | <LOD | 20 | 361 | 357 | 119 | 5 | 112 | 258 | 38 | 309 | 24 | 11 | <LOD | 76 | 48 | 116 | 915 | 42 | <LOD | 10 | 1296 | 22 | <LOD | 4 | 731107
5 [Col] | 2023-06-14 | NIST 2711a | std | rep 5 | 15557 | 67096 | 262799 | 230 | 186 | 25287 | 22628 | 3107 | <LOD | 138 | 660 | 28728 | 200 | 25 | 381 | 395 | 141 | 3 | 120 | 266 | 25 | 329 | 21 | 12 | 24 | 50 | 45 | 97 | 1052 | 18 | <LOD | 8 | 1356 | <LOD | 11 | <LOD | 728699
6 [Col] | 2023-06-14 | NIST 679 | std | rep 1 | 11288 | 93410 | 221930 | <LOD | <LOD | 24021 | 1455 | 5406 | 270 | 119 | 1458 | 93463 | <LOD | 52 | 373 | 112 | 21 | 3 | 183 | 47 | 45 | 171 | 22 | 18 | 28 | 25 | 56 | 73 | <LOD | 41 | <LOD | 16 | 15 | <LOD | <LOD | <LOD | 700069
7 [Col] | 2023-06-14 | NIST 679 | std | rep 2 | 9769 | 91158 | 219610 | <LOD | <LOD | 24175 | 1481 | 5849 | 136 | 147 | 1546 | 93895 | 116 | 47 | 403 | 120 | 25 | 4 | 197 | 50 | 42 | 162 | 23 | 9 | 15 | <LOD | <LOD | 36 | <LOD | 34 | 4 | <LOD | 11 | <LOD | 37 | <LOD | 704703
8 [Col] | 2023-06-14 | NIST 679 | std | rep 3 | 11661 | 94568 | 222879 | <LOD | <LOD | 24054 | 1374 | 4994 | 155 | <LOD | 1577 | 91039 | <LOD | 83 | 428 | 144 | 20 | 3 | 188 | 43 | 35 | 157 | 20 | 9 | <LOD | 18 | 70 | 87 | 767 | <LOD | 9 | 8 | 16 | <LOD | 23 | <LOD | 700031
9 [Col] | 2023-06-14 | NIST 679 | std | rep 4 | 11180 | 91140 | 216874 | <LOD | <LOD | 23287 | 1344 | 5497 | 261 | 112 | 1763 | 90819 | <LOD | 72 | 442 | 122 | 16 | 7 | 182 | 54 | 42 | 162 | 26 | 11 | 26 | 35 | 61 | 88 | 951 | 20 | <LOD | 19 | 15 | <LOD | 13 | <LOD | 708327
10 [Col] | 2023-06-14 | NIST 679 | std | rep 5 | 11945 | 92074 | 221004 | <LOD | <LOD | 23542 | 1402 | 5109 | 186 | 126 | 1513 | 91839 | <LOD | 81 | 386 | 122 | 23 | <LOD | 197 | 48 | 38 | 162 | 27 | 16 | 74 | 47 | 71 | 66 | 1144 | 44 | <LOD | <LOD | 22 | <LOD | 13 | <LOD | 702200
11 [Col] | 2023-06-14 | SARM 69 | std | rep 1 | 12560 | 77115 | 262716 | 299 | <LOD | 17200 | 13598 | 4180 | 134 | 238 | 1031 | 51261 | 201 | 68 | 320 | 61 | 5 | 5 | 82 | 105 | 35 | 284 | 6 | 14 | 18 | 40 | 52 | 53 | 1816 | 33 | <LOD | <LOD | 21 | 20 | <LOD | <LOD | 712065
12 [Col] | 2023-06-14 | SARM 69 | std | rep 2 | 13010 | 76707 | 261540 | 226 | <LOD | 17259 | 13685 | 4271 | 96 | 146 | 957 | 50681 | 150 | 73 | 292 | 69 | 14 | 5 | 81 | 116 | 35 | 303 | 9 | 13 | <LOD | 14 | 37 | 58 | 1871 | <LOD | <LOD | <LOD | 11 | 19 | <LOD | <LOD | 713691
13 [Col] | 2023-06-14 | SARM 69 | std | rep 3 | 13069 | 77628 | 263263 | 344 | <LOD | 17570 | 13832 | 4381 | 159 | 116 | 1026 | 51397 | 99 | 37 | 316 | 61 | 5 | 3 | 77 | 98 | 35 | 303 | 15 | 6 | 32 | 37 | 46 | 39 | 817 | 30 | <LOD | <LOD | 14 | <LOD | 15 | <LOD | 711070
14 [Col] | 2023-06-14 | SARM 69 | std | rep 4 | 12900 | 77463 | 264651 | 412 | <LOD | 17490 | 14606 | 4927 | 155 | 335 | 960 | 51507 | 215 | 72 | 329 | 78 | 7 | 4 | 80 | 109 | 34 | 296 | 4 | 6 | 29 | 18 | 85 | 104 | 829 | 46 | <LOD | <LOD | 18 | 20 | <LOD | <LOD | 708457
15 [Col] | 2023-06-14 | SARM 69 | std | rep 5 | 13411 | 77436 | 261143 | 287 | <LOD | 17471 | 14485 | 4143 | 102 | 213 | 862 | 50336 | 212 | 71 | 316 | 67 | 9 | 4 | 77 | 100 | 28 | 246 | 12 | 5 | <LOD | 43 | <LOD | 36 | 1051 | <LOD | 4 | 6 | 14 | 36 | <LOD | <LOD | 713671
1 [Col] | 2023-06-15 | NIST 2711a | std | rep 1 | 14406 | 70601 | 266886 | 194 | 101 | 25791 | 23140 | 3702 | 79 | 118 | 697 | 29017 | 66 | 20 | 271 | 395 | 161 | 5 | 121 | 258 | 28 | 323 | 18 | 6 | 16 | 87 | 51 | 60 | 737 | 38 | <LOD | <LOD | 1361 | <LOD | <LOD | <LOD | 722563
2 [Col] | 2023-06-15 | NIST 2711a | std | rep 2 | 14507 | 69332 | 262826 | 204 | 170 | 25689 | 23268 | 3291 | 127 | 115 | 794 | 29907 | 95 | 28 | 316 | 381 | 143 | 4 | 131 | 258 | 30 | 342 | 22 | 16 | 20 | 97 | 74 | 102 | 954 | 25 | 7 | 9 | 1372 | <LOD | 8 | <LOD | 726077
3 [Col] | 2023-06-15 | NIST 2711a | std | rep 3 | 15879 | 69400 | 264460 | 161 | 130 | 24858 | 23022 | 3356 | <LOD | 73 | 697 | 28784 | 66 | 32 | 261 | 377 | 100 | 4 | 116 | 266 | 36 | 306 | 26 | 12 | <LOD | 73 | 56 | 64 | 1780 | 34 | <LOD | 9 | 1338 | <LOD | 10 | <LOD | 724445
4 [Col] | 2023-06-15 | NIST 2711a | std | rep 4 | 15055 | 69317 | 263373 | 123 | 128 | 25377 | 22040 | 2987 | <LOD | 95 | 779 | 28625 | <LOD | 37 | 297 | 370 | 98 | 2 | 122 | 261 | 34 | 264 | 18 | 11 | 25 | 90 | 47 | 91 | 1027 | 21 | 4 | 6 | 1349 | <LOD | 9 | <LOD | 728047
5 [Col] | 2023-06-15 | NIST 2711a | std | rep 5 | 14797 | 69450 | 264869 | 221 | 127 | 26161 | 23097 | 3110 | 85 | 57 | 632 | 29923 | <LOD | 34 | 248 | 411 | 148 | 6 | 117 | 268 | 32 | 311 | 9 | 6 | 24 | 77 | 51 | 91 | 1315 | 25 | 6 | <LOD | 1373 | <LOD | 20 | <LOD | 723916
6 [Col] | 2023-06-15 | NIST 679 | std | rep 1 | 11574 | 90355 | 216641 | <LOD | <LOD | 23352 | 1467 | 5431 | 81 | 100 | 1621 | 89805 | 326 | 51 | 219 | 114 | 23 | 4 | 197 | 55 | 39 | 163 | 21 | 15 | 24 | 27 | 61 | 138 | 961 | 32 | <LOD | <LOD | 4 | 20 | <LOD | <LOD | 709885
7 [Col] | 2023-06-15 | NIST 679 | std | rep 2 | 9769 | 89292 | 216964 | <LOD | <LOD | 23585 | 1423 | 5797 | 167 | 131 | 1823 | 92308 | <LOD | 79 | 260 | 113 | 24 | 6 | 196 | 55 | 39 | 155 | 21 | 23 | <LOD | 44 | 74 | 45 | 608 | 46 | <LOD | <LOD | 9 | <LOD | 24 | <LOD | 709846
8 [Col] | 2023-06-15 | NIST 679 | std | rep 3 | 9769 | 94512 | 223310 | <LOD | <LOD | 24433 | 1459 | 5536 | 155 | 126 | 1674 | 94985 | 196 | 63 | 231 | 120 | 24 | 2 | 206 | 57 | 47 | 179 | 22 | 20 | 28 | 29 | <LOD | 78 | 748 | 48 | <LOD | <LOD | 13 | <LOD | 30 | <LOD | 696835
9 [Col] | 2023-06-15 | NIST 679 | std | rep 4 | 11058 | 92185 | 221121 | <LOD | <LOD | 24018 | 1449 | 5470 | 283 | 50 | 1484 | 90475 | 177 | 49 | 244 | 121 | 19 | 3 | 189 | 49 | 48 | 157 | 30 | 9 | 15 | 34 | 105 | 114 | 476 | 41 | <LOD | <LOD | 25 | <LOD | 16 | 8 | 704375
10 [Col] | 2023-06-15 | NIST 679 | std | rep 5 | 12225 | 91954 | 221134 | <LOD | <LOD | 23924 | 1403 | 5982 | <LOD | 82 | 1569 | 92175 | 162 | 74 | 257 | 132 | 23 | 7 | 187 | 56 | 38 | 165 | 35 | 24 | 21 | 41 | 50 | 112 | 827 | <LOD | <LOD | 13 | 8 | <LOD | 9 | <LOD | 700865
11 [Col] | 2023-06-15 | SARM 69 | std | rep 1 | 13995 | 78039 | 266307 | 239 | <LOD | 17535 | 14623 | 4409 | 270 | 200 | 961 | 51043 | 93 | 65 | 177 | 62 | 7 | 7 | 73 | 107 | 35 | 361 | 10 | 11 | 30 | 37 | 43 | 36 | <LOD | <LOD | 6 | <LOD | 17 | 23 | <LOD | <LOD | 707571
12 [Col] | 2023-06-15 | SARM 69 | std | rep 2 | 13442 | 78961 | 266243 | 320 | <LOD | 17688 | 13807 | 4567 | 124 | 193 | 1006 | 50732 | 63 | 51 | 179 | 72 | 6 | 4 | 76 | 106 | 35 | 307 | 11 | 15 | <LOD | 24 | 54 | 71 | 529 | 26 | <LOD | 10 | 15 | <LOD | 13 | 4 | 707672
13 [Col] | 2023-06-15 | SARM 69 | std | rep 3 | 13222 | 78058 | 265269 | 252 | <LOD | 17463 | 14277 | 4467 | 78 | 261 | 1079 | 50697 | 193 | 44 | 210 | 57 | 10 | 3 | 80 | 113 | 37 | 262 | 8 | 7 | 13 | 52 | 49 | 35 | 1532 | 36 | <LOD | 8 | 12 | <LOD | 16 | <LOD | 708383
14 [Col] | 2023-06-15 | SARM 69 | std | rep 4 | 12711 | 79497 | 269757 | 331 | <LOD | 17918 | 14094 | 4418 | 167 | 241 | 943 | 51999 | 136 | 49 | 200 | 79 | 16 | 4 | 76 | 106 | 29 | 265 | 4 | <LOD | <LOD | 15 | 54 | 89 | 617 | <LOD | <LOD | 5 | 8 | <LOD | <LOD | <LOD | 703313
15 [Col] | 2023-06-15 | SARM 69 | std | rep 5 | 13255 | 77463 | 264568 | 271 | <LOD | 17723 | 14086 | 4653 | 124 | 200 | 1019 | 52020 | 238 | 77 | 207 | 74 | 4 | 3 | 85 | 106 | 33 | 276 | 8 | 11 | <LOD | <LOD | 54 | 75 | 1028 | 17 | <LOD | 10 | 16 | <LOD | 29 | <LOD | 708366
1 [Col] | 2023-06-16 | NIST 2711a | std | rep 1 | 14384 | 66551 | 259347 | 82 | 173 | 25044 | 22069 | 3362 | 137 | 84 | 629 | 28466 | 77 | 27 | 303 | 383 | 126 | 4 | 122 | 253 | 31 | 285 | 18 | 13 | <LOD | 81 | 36 | 79 | 1296 | <LOD | <LOD | 8 | 1355 | <LOD | 13 | <LOD | 734332
2 [Col] | 2023-06-16 | NIST 2711a | std | rep 2 | 15319 | 68885 | 261955 | 108 | 109 | 25497 | 23231 | 3742 | <LOD | 84 | 706 | 29432 | 75 | 27 | 327 | 411 | 146 | <LOD | 127 | 263 | 35 | 347 | 26 | <LOD | 33 | 77 | 36 | 85 | 1363 | 43 | <LOD | 13 | 1380 | <LOD | <LOD | 8 | 726419
3 [Col] | 2023-06-16 | NIST 2711a | std | rep 3 | 14470 | 68749 | 258949 | 94 | 100 | 25251 | 22103 | 3313 | 139 | 100 | 673 | 28061 | <LOD | 29 | 251 | 372 | 89 | 6 | 125 | 255 | 31 | 283 | 22 | 7 | 30 | 86 | 50 | 101 | <LOD | <LOD | 5 | 7 | 1334 | 24 | <LOD | <LOD | 734604
4 [Col] | 2023-06-16 | NIST 2711a | std | rep 4 | 15494 | 69119 | 263320 | 175 | 229 | 25460 | 23137 | 4278 | 196 | 104 | 668 | 30162 | <LOD | 24 | 320 | 395 | 176 | <LOD | 123 | 268 | 33 | 337 | 24 | 10 | <LOD | 106 | 59 | 80 | 614 | 41 | <LOD | <LOD | 1356 | <LOD | 13 | <LOD | 724037
5 [Col] | 2023-06-16 | NIST 2711a | std | rep 5 | 14739 | 69461 | 261636 | 166 | 117 | 25422 | 22445 | 3150 | <LOD | 69 | 565 | 29808 | 75 | 32 | 263 | 391 | 173 | 3 | 111 | 268 | 32 | 307 | 18 | 7 | 22 | 97 | 50 | 117 | 1332 | <LOD | <LOD | 16 | 1301 | <LOD | 20 | <LOD | 728028
6 [Col] | 2023-06-16 | NIST 679 | std | rep 1 | 12730 | 94004 | 223819 | <LOD | <LOD | 24236 | 1413 | 6303 | 345 | 45 | 1706 | 92541 | 184 | 66 | 264 | 142 | 23 | 9 | 193 | 53 | 38 | 177 | 24 | 16 | <LOD | 48 | 37 | 35 | 577 | 54 | <LOD | <LOD | 13 | 30 | <LOD | 6 | 695096
7 [Col] | 2023-06-16 | NIST 679 | std | rep 2 | 10938 | 92600 | 221779 | <LOD | <LOD | 24070 | 1463 | 5217 | 322 | 169 | 1374 | 89439 | 152 | 68 | 229 | 167 | 26 | 4 | 184 | 50 | 40 | 154 | 25 | 8 | 18 | 42 | 82 | 36 | <LOD | <LOD | 4 | <LOD | 19 | <LOD | 12 | 6 | 705402
8 [Col] | 2023-06-16 | NIST 679 | std | rep 3 | 9769 | 95182 | 221733 | <LOD | <LOD | 24061 | 1476 | 5487 | 105 | 124 | 1617 | 93442 | 145 | 42 | 237 | 127 | 15 | 7 | 199 | 62 | 34 | 169 | 24 | 17 | 24 | 17 | 29 | 50 | 1247 | 20 | <LOD | 8 | 21 | <LOD | <LOD | <LOD | 699339
9 [Col] | 2023-06-16 | NIST 679 | std | rep 4 | 13102 | 91256 | 221472 | <LOD | <LOD | 24136 | 1492 | 5958 | 137 | 123 | 1711 | 93493 | <LOD | 81 | 227 | 110 | 28 | 4 | 195 | 51 | 38 | 169 | 22 | 5 | 22 | 37 | 87 | 57 | 539 | 16 | <LOD | <LOD | 8 | <LOD | 19 | <LOD | 698803
10 [Col] | 2023-06-16 | NIST 679 | std | rep 5 | 11808 | 92052 | 220513 | <LOD | <LOD | 24044 | 1473 | 5214 | 148 | 41 | 1619 | 93357 | <LOD | 45 | 207 | 124 | 24 | 3 | 185 | 40 | 30 | 160 | 24 | 10 | <LOD | 57 | 71 | 93 | 666 | 18 | <LOD | <LOD | 4 | <LOD | 17 | 6 | 701673
11 [Col] | 2023-06-16 | SARM 69 | std | rep 1 | 12628 | 78743 | 267030 | 302 | <LOD | 17250 | 14035 | 4809 | 99 | 228 | 856 | 50693 | <LOD | 55 | 194 | 59 | 13 | 3 | 78 | 99 | 32 | 259 | 9 | <LOD | 44 | 61 | 35 | 27 | 570 | 33 | <LOD | 4 | 12 | <LOD | <LOD | <LOD | 708200
12 [Col] | 2023-06-16 | SARM 69 | std | rep 2 | 13066 | 79366 | 269789 | 392 | <LOD | 17653 | 14746 | 4724 | 108 | 184 | 964 | 52161 | 255 | 61 | 221 | 73 | 6 | 7 | 75 | 111 | 34 | 301 | 7 | 7 | 14 | 35 | 59 | 69 | 402 | <LOD | <LOD | <LOD | 21 | <LOD | 22 | 4 | 702234
13 [Col] | 2023-06-16 | SARM 69 | std | rep 3 | 11538 | 78915 | 266501 | 320 | <LOD | 17407 | 14066 | 4503 | 249 | 293 | 1000 | 51206 | 71 | 73 | 216 | 73 | 5 | 2 | 79 | 109 | 35 | 299 | 15 | 11 | 12 | 25 | 36 | 90 | 1128 | <LOD | 8 | 8 | 17 | 42 | <LOD | <LOD | 708378
14 [Col] | 2023-06-16 | SARM 69 | std | rep 4 | 13429 | 78984 | 266633 | 275 | <LOD | 17858 | 14356 | 4909 | 178 | 198 | 1018 | 51021 | 273 | 56 | 235 | 78 | 5 | 3 | 76 | 104 | 29 | 332 | 8 | 14 | <LOD | <LOD | 64 | 67 | <LOD | 25 | <LOD | 9 | 15 | <LOD | 8 | <LOD | 706452
15 [Col] | 2023-06-16 | SARM 69 | std | rep 5 | 12648 | 78607 | 269833 | 296 | <LOD | 17704 | 14409 | 4361 | 155 | 161 | 992 | 50157 | 186 | 53 | 231 | 79 | 4 | 8 | 77 | 107 | 27 | 309 | 8 | <LOD | 17 | 27 | 57 | 79 | 881 | <LOD | <LOD | <LOD | 17 | <LOD | 11 | <LOD | 705505
2-RP1 [Col] | 2023-07-18 | NIST 2711a | std | rep 1 | 12475.8 | 65610.4 | 256441.7 | <LOD | <LOD | 23905.4 | 21743.3 | 2681 | 113.9 | 34.4 | 614.2 | 27652 | 161.9 | 26.4 | 172.2 | 364.3 | 118.4 | 2 | 114.2 | 257.1 | 36.9 | 267.9 | 20.6 | 10.8 | 22.5 | 93.7 | 77.3 | 117.2 | 1038.9 | <LOD | <LOD | 7.2 | 1274.1 | <LOD | 11.3 | <LOD | 743061.1
3 [Col] | 2023-07-18 | NIST 2711a | std | rep 2 | 9769 | 63834.3 | 247201 | <LOD | <LOD | 24005.5 | 21803.1 | 2925.2 | <LOD | 105.7 | 630.4 | 27469.6 | 159.7 | 21.1 | 218.8 | 391.1 | 147.3 | <LOD | 125.3 | 252.4 | 28.5 | 325.1 | 16.6 | 8.6 | 37.6 | 82.7 | 34.9 | 77.4 | 1228 | <LOD | <LOD | <LOD | 1316.6 | 32.6 | <LOD | <LOD | 755655.9
4 [Col] | 2023-07-18 | NIST 2711a | std | rep 3 | 11619.1 | 66900.5 | 261294.4 | <LOD | <LOD | 25475.5 | 22424.2 | 3200.3 | 73.1 | 92.7 | 555.7 | 28425.1 | <LOD | 14.4 | 202.6 | 399.1 | 107.7 | 3.5 | 113.2 | 253.3 | 30.6 | 277.5 | 25 | 9.1 | 13.7 | 61.3 | 31.8 | 93.2 | 847 | <LOD | <LOD | 4.4 | 1335.2 | 47.6 | <LOD | 4.1 | 736244.1
5 [Col] | 2023-07-18 | NIST 2711a | std | rep 4 | 9769 | 64800.2 | 250249 | <LOD | <LOD | 23505.5 | 22132.9 | 3275.4 | 258.3 | <LOD | 741.7 | 28430.5 | 151.5 | 24.6 | 196.6 | 356 | 154.3 | 2.9 | 118.4 | 257.6 | 35.6 | 292.9 | 21.7 | 20.6 | 12.9 | 86 | 29.1 | 114.6 | 1539.7 | 16.4 | <LOD | 14.5 | 1302.8 | <LOD | <LOD | <LOD | 750336.2
6 [Col] | 2023-07-18 | NIST 2711a | std | rep 5 | 11692.3 | 68982.8 | 265943.1 | <LOD | <LOD | 25462.1 | 22701.2 | 3740.4 | <LOD | <LOD | 688 | 28966.6 | 123.5 | 23 | 193.7 | 385.7 | 142.2 | 3.3 | 111.4 | 261.7 | 29.3 | 319.5 | 18.3 | 16.2 | <LOD | 86.9 | 59.6 | 108.1 | 878.3 | 41.8 | <LOD | 7.1 | 1285.1 | <LOD | 6.8 | 4.2 | 728763
7 [Col] | 2023-07-18 | NIST 679 | std | rep 1 | 9769 | 85224.5 | 214541.3 | <LOD | <LOD | 23038.6 | 1523.1 | 5225 | 111.3 | 91.9 | 1529.5 | 88109.3 | 215.2 | 49.5 | 106.2 | 142.1 | 17 | 4.4 | 187.9 | 49.5 | 44.6 | 169.3 | 18.8 | 9.4 | 45.3 | 39.9 | 59.1 | 80.2 | 1139.4 | 17.7 | <LOD | <LOD | 18.5 | <LOD | 28.1 | <LOD | 720117.7
8 [Col] | 2023-07-18 | NIST 679 | std | rep 2 | 9769 | 84741.8 | 211765.6 | <LOD | <LOD | 22495.5 | 1526.7 | 5254.3 | <LOD | 52 | 1522 | 89422.6 | 248.9 | 50.3 | 118.9 | 112.2 | 13.6 | 4.7 | 188.9 | 40.6 | 30.7 | 161.1 | 22.7 | 7.7 | 15.1 | 65.2 | 63.9 | 84 | 618.5 | 15.8 | <LOD | <LOD | 20.4 | <LOD | 27.3 | <LOD | 722720.5
9 [Col] | 2023-07-18 | NIST 679 | std | rep 3 | 9769 | 89634.3 | 219125.9 | <LOD | <LOD | 23663.7 | 1498.9 | 4804.2 | <LOD | 50.8 | 1587.5 | 91516.9 | 253.6 | 24.7 | 99.5 | 121.4 | 25.7 | <LOD | 194.8 | 49.6 | 43.7 | 168.9 | 35 | 13.8 | 29.4 | 69.7 | 27.7 | 45.7 | 1315.5 | 39.7 | <LOD | <LOD | 9.3 | <LOD | 43.7 | <LOD | 709055.9
10 [Col] | 2023-07-18 | NIST 679 | std | rep 4 | 9769 | 92373.8 | 219717.3 | <LOD | <LOD | 23560.6 | 1447.4 | 5343.3 | 73 | 95.7 | 1740.5 | 88720.2 | 261.3 | 27.7 | 104.2 | 115.2 | 28.4 | 6.1 | 192.7 | 44.6 | 39 | 150.5 | 25.8 | 21.3 | 21.8 | 24.8 | 74.2 | 45.2 | 973.9 | 42.7 | <LOD | <LOD | 4 | 21.2 | 8.8 | <LOD | 708796
11 [Col] | 2023-07-18 | NIST 679 | std | rep 5 | 9769 | 73652.7 | 183996 | <LOD | <LOD | 16881.1 | 1122.8 | 3561.2 | 106.2 | 56.4 | 995 | 56927.3 | 159.4 | 16.9 | 72.6 | 76 | 16.9 | 2.9 | 108.1 | 6.3 | 27.2 | 87.3 | 11.6 | 13.9 | <LOD | 13 | 21.3 | 65.2 | 506 | 17.3 | <LOD | <LOD | 4 | 22.3 | <LOD | <LOD | 796111
12 [Col] | 2023-07-18 | SARM 69 | std | rep 1 | 11453.4 | 76044 | 263795.4 | 152.8 | <LOD | 17074.8 | 13967.1 | 4780.2 | 151.8 | 200.4 | 954.4 | 51304.9 | 135.3 | 45.8 | 130.7 | 63.3 | 9.4 | 4.8 | 68.3 | 95.2 | 35.3 | 288.1 | 13.3 | 4.5 | 18.5 | 31.9 | 70.9 | 49 | <LOD | 13.5 | <LOD | 8.1 | 12.7 | <LOD | <LOD | 4.6 | 714734.4
13 [Col] | 2023-07-18 | SARM 69 | std | rep 2 | 11444.4 | 77798.8 | 265877.3 | 155.6 | <LOD | 17301.9 | 14126.7 | 4518.7 | 144.5 | 173.8 | 958.8 | 51046 | 344.9 | 39.5 | 123.7 | 65.9 | 14.4 | 3.9 | 71.5 | 95.1 | 30 | 263.7 | 8.6 | 7.4 | 34.4 | 44.3 | 42 | 33.7 | 1241.5 | 31.2 | <LOD | <LOD | 10.6 | <LOD | <LOD | 8.9 | 710375.1
14 [Col] | 2023-07-18 | SARM 69 | std | rep 3 | 11980.9 | 74935.1 | 264070.2 | 90.3 | <LOD | 16609.1 | 13955.4 | 4525.7 | 104.1 | 189.2 | 923.5 | 50425.8 | 70.6 | 55.6 | 127 | 56.8 | 9.5 | 4.1 | 71.5 | 98.5 | 31.7 | 290.5 | 3.5 | 23.6 | 28 | 37.9 | 47.1 | 43.8 | 1894.9 | <LOD | <LOD | <LOD | 13.5 | <LOD | 9.2 | <LOD | 714508.1
15 [Col] | 2023-07-18 | SARM 69 | std | rep 4 | 11664.2 | 77358.8 | 268069 | 157.3 | <LOD | 17333.2 | 14237.9 | 4495.1 | 171.7 | 115.6 | 962.5 | 50623 | 158.3 | 58 | 98.4 | 58.3 | 4.3 | 4.4 | 77.2 | 107.5 | 29.5 | 264 | 10 | 9.7 | <LOD | 36.9 | 51.1 | 47.1 | 456.3 | 57.9 | <LOD | <LOD | 12.8 | <LOD | 15.6 | <LOD | 709786.8
16 [Col] | 2023-07-18 | SARM 69 | std | rep 5 | 12951 | 78952.1 | 265823 | 142.6 | <LOD | 17384.4 | 13846.8 | 4592.8 | 124.3 | 197.5 | 845.2 | 51522.2 | 121.5 | 59 | 89.9 | 65.3 | 8.5 | <LOD | 76.8 | 101.7 | 30.5 | 419 | 9.4 | 11 | <LOD | 17.7 | 80.4 | 118 | 1149.7 | 37.1 | <LOD | <LOD | 16.8 | 18.6 | <LOD | <LOD | 707548.8
2 [Col] | 2023-07-28 | NIST 2711a | std | rep 1 | 14363.5 | 69069.9 | 265403.5 | <LOD | <LOD | 26057 | 23132.9 | 3397.4 | 85.4 | 80 | 699.2 | 28811.4 | 66.2 | 26.4 | 172 | 370.6 | 148.8 | 2.9 | 119.2 | 263.8 | 30.8 | 266.4 | 21.5 | 13.1 | 14.3 | 93.1 | 21.6 | 67.3 | 1066.1 | <LOD | <LOD | 19 | 1298.7 | <LOD | 19 | <LOD | 725815
3 [Col] | 2023-07-28 | NIST 2711a | std | rep 2 | 11917.4 | 67136.1 | 259771.4 | <LOD | <LOD | 25097.9 | 23342.1 | 3270.3 | 84.2 | <LOD | 723.1 | 28588.1 | 120.4 | 18.8 | 208 | 397 | 89 | 7.6 | 125.1 | 262.4 | 25.8 | 329.9 | 25.8 | 5.4 | 29.1 | 106.9 | 44.7 | 104.2 | 1057.6 | 34 | <LOD | 8.5 | 1412.9 | <LOD | 11.6 | <LOD | 735882.3
4 [Col] | 2023-07-28 | NIST 2711a | std | rep 3 | 12926.4 | 69693.2 | 269941.9 | <LOD | <LOD | 26351.3 | 23012.4 | 3636.4 | 132.1 | 47.4 | 690.1 | 29321.3 | 138.5 | 19.8 | 205 | 398.9 | 130.2 | 2.4 | 121.7 | 268.1 | 38.2 | 252.8 | 24.2 | 15.6 | 17.9 | 80 | 66.8 | 106.2 | 916.8 | <LOD | <LOD | 10.1 | 1395.7 | <LOD | 10 | 7.4 | 721803
5 [Col] | 2023-07-28 | NIST 2711a | std | rep 4 | 13501.8 | 67772.7 | 260656.6 | <LOD | <LOD | 25619 | 23480.3 | 3790.2 | 210.3 | 41.9 | 655.6 | 28837.4 | 99.1 | 27.7 | 219.2 | 402.4 | 133.6 | <LOD | 125.4 | 258.1 | 28.8 | 307.1 | 21.8 | 7 | 22.2 | 82.6 | 40 | 98.6 | 792.8 | 27.1 | <LOD | 9 | 1343.6 | <LOD | 12.4 | <LOD | 731790.6
6 [Col] | 2023-07-28 | NIST 2711a | std | rep 5 | 14297.8 | 70706.6 | 269225.1 | <LOD | <LOD | 25800.2 | 23144.4 | 3501.6 | <LOD | <LOD | 629.3 | 29407.9 | 106.2 | 9.1 | 176.7 | 386.9 | 120.4 | 4.1 | 126.1 | 267 | 32.1 | 285.1 | 15 | 6.1 | 27.9 | 76.7 | 69.2 | 80 | 1307.3 | 44.7 | <LOD | <LOD | 1331.2 | <LOD | 12 | <LOD | 720369.8
7 [Col] | 2023-07-28 | NIST 679 | std | rep 1 | 9769 | 91404.6 | 221429 | <LOD | <LOD | 24130 | 1562.5 | 5685.5 | 105.7 | 80.8 | 1585.7 | 93651.3 | 180.5 | 49.4 | 80.1 | 145 | 26.5 | 7.6 | 191.7 | 55.5 | 39.8 | 168.5 | 16.1 | 10.7 | 19.1 | 27.3 | 65.2 | 31.7 | 716 | <LOD | 4.9 | <LOD | 11.8 | <LOD | 22.2 | 4.3 | 702739.1
8 [Col] | 2023-07-28 | NIST 679 | std | rep 2 | 9769 | 86250.2 | 213640.4 | <LOD | <LOD | 23174.9 | 1396.9 | 5418.8 | 79.3 | 43.3 | 1684.3 | 89590.2 | 280.3 | 49.3 | 55 | 125.1 | 28.9 | 2.9 | 187.7 | 51.9 | 39.2 | 154 | 25.9 | <LOD | <LOD | 20.4 | 46.8 | 91.6 | 1052 | <LOD | 6.3 | <LOD | 4 | <LOD | 20 | <LOD | 718582
9 [Col] | 2023-07-28 | NIST 679 | std | rep 3 | 9769 | 80961.1 | 202047.3 | <LOD | <LOD | 21293 | 1325.3 | 5115.5 | 236.2 | 129.3 | 1350 | 79404.3 | <LOD | 47.2 | 72.1 | 119.5 | 27 | 2.5 | 160.4 | 32.5 | 31.4 | 123.9 | 11.5 | 4.3 | <LOD | 60.4 | 52.2 | 59.5 | 339.5 | <LOD | <LOD | <LOD | 4 | <LOD | 18.4 | <LOD | 746288
10 [Col] | 2023-07-28 | NIST 679 | std | rep 4 | 9769 | 82882.2 | 207972.3 | <LOD | <LOD | 22277.3 | 1384.9 | 5431.9 | 174.3 | 39.1 | 1608.3 | 86798.3 | <LOD | 69.6 | 44 | 107.1 | 18.9 | 4.7 | 166.9 | 35 | 36.8 | 145 | 21.5 | 9.4 | <LOD | 22.5 | 82.5 | 105 | 897 | 32.9 | <LOD | <LOD | 8.7 | <LOD | 14.6 | <LOD | 730172
11 [Col] | 2023-07-28 | NIST 679 | std | rep 5 | 9769 | 92709.2 | 220575.4 | <LOD | <LOD | 23891.6 | 1461.4 | 5831.4 | 210.6 | 99.2 | 1676.2 | 91518.3 | 167.1 | 53.9 | 68.4 | 114.1 | 31.6 | 7.3 | 191.5 | 46.4 | 48.4 | 172.8 | 16.9 | 5.9 | 20.4 | 18.8 | 82.2 | 145.3 | <LOD | 37.6 | <LOD | <LOD | 9.2 | <LOD | 14.3 | <LOD | 705104.9
12 [Col] | 2023-07-28 | SARM 69 | std | rep 1 | 13925.2 | 78545.8 | 271282.9 | 129.3 | <LOD | 17273.4 | 14552.1 | 4879.2 | 88 | 217.2 | 951.6 | 50217.9 | 112.7 | 62 | 52.9 | 63.9 | 6.5 | 3.6 | 72 | 93.4 | 29.9 | 268.8 | 6.5 | 9 | <LOD | 42.4 | 40.7 | 81.7 | 853.5 | 27.3 | <LOD | <LOD | 18.2 | <LOD | <LOD | <LOD | 702823.9
13 [Col] | 2023-07-28 | SARM 69 | std | rep 2 | 13011.4 | 77452.3 | 273544.7 | 141.8 | <LOD | 17661.7 | 14765 | 4989.5 | 189.3 | 189.3 | 833.8 | 51730.5 | <LOD | 45.1 | 70.7 | 59.8 | 9.5 | 4.1 | 75.7 | 99.2 | 33.5 | 292.4 | 9.1 | 6.7 | 22.9 | 26.2 | 46.6 | 151.1 | 716.8 | 12.9 | <LOD | <LOD | 9.2 | 21.8 | <LOD | <LOD | 700829.5
14 [Col] | 2023-07-28 | SARM 69 | std | rep 3 | 12523.9 | 78836.5 | 268195.7 | 152.2 | <LOD | 17804.8 | 14523.1 | 4495 | 162.5 | 197.9 | 900.2 | 51004.9 | 164.2 | 42.9 | 95.6 | 81.8 | 5.7 | 3.9 | 81 | 108.4 | 30.4 | 250 | 5.5 | 14.8 | 32.5 | 35.8 | 48.4 | 76.9 | 925 | <LOD | <LOD | <LOD | 18.5 | <LOD | 9.5 | <LOD | 706193.1
15 [Col] | 2023-07-28 | SARM 69 | std | rep 4 | 13389.7 | 79767.8 | 272812.3 | 178.1 | <LOD | 17688.9 | 14853.1 | 4889.1 | 176.2 | 199.8 | 1035.5 | 51802.3 | <LOD | 44.1 | 91.4 | 69.4 | 5.6 | 6.2 | 77.7 | 106.6 | 29.4 | 282.4 | 7.5 | 4.8 | <LOD | 18.6 | 48.4 | 104.2 | 1069.6 | 13.5 | <LOD | 4.4 | 17.2 | <LOD | <LOD | <LOD | 698691
16 [Col] | 2023-07-28 | SARM 69 | std | rep 5 | 13046.8 | 79526.7 | 275532.9 | 156.4 | <LOD | 17931.8 | 14781.8 | 5100.6 | 290.1 | 221.9 | 955.9 | 51382.4 | 258 | 64.9 | 90.3 | 68.6 | 13.1 | 4.5 | 79.2 | 103.7 | 30.1 | 305.5 | 9.5 | 5.2 | 30.4 | 75.2 | 82.6 | 110.7 | 514.6 | 17.6 | <LOD | <LOD | 11 | <LOD | 18 | 7.1 | 696930
1 [Col] | 2023-07-31 | NIST 2711a | std | rep 1 | 12707.1 | 69973.2 | 270704 | <LOD | <LOD | 26262.8 | 24432.2 | 3213.7 | 80.8 | 31.7 | 607.4 | 29322.6 | 134.5 | 27.2 | 210.5 | 408.1 | 145.8 | 3.6 | 125 | 273.2 | 30.5 | 276.6 | 22.4 | 8 | 36 | 106.8 | <LOD | <LOD | 979.6 | <LOD | <LOD | 18 | 1381.1 | <LOD | 19.5 | <LOD | 720924.6
2 [Col] | 2023-07-31 | NIST 2711a | std | rep 2 | 9769 | 68773.2 | 265099.5 | <LOD | <LOD | 25923.5 | 23518 | 3431.5 | 149.9 | <LOD | 704.5 | 29575.7 | <LOD | 17.2 | 169.5 | 421.6 | 156.4 | 3.6 | 121.5 | 255.9 | 31.8 | 305.4 | 21.9 | 17.3 | 40.6 | 91.5 | <LOD | 37.8 | <LOD | 21.3 | <LOD | 10.8 | 1371.3 | <LOD | 16.5 | 7.4 | 731682.8
3 [Col] | 2023-07-31 | NIST 2711a | std | rep 3 | 13621.3 | 69320.5 | 266486.9 | 43.1 | <LOD | 25504.7 | 23554.3 | 3314.9 | <LOD | 35.3 | 777.2 | 29070 | 146.7 | 18.9 | 193.1 | 377.6 | 136.7 | 4 | 120.1 | 249.9 | 37.8 | 341.5 | 28.1 | 11 | <LOD | 75.8 | 42.9 | 73.4 | 963.8 | 35.9 | <LOD | 11.4 | 1349.7 | <LOD | <LOD | <LOD | 725256.9
4 [Col] | 2023-07-31 | NIST 2711a | std | rep 4 | 14380.6 | 72043.2 | 269929.3 | 61.3 | <LOD | 26112 | 23385 | 3464.4 | 219.1 | 49.6 | 568.4 | 29471.8 | 52.8 | 3 | 187.1 | 375.7 | 142.2 | 3 | 122.3 | 261 | 22.2 | 285.1 | 20.9 | 4.5 | 17.1 | 102.7 | 75.3 | 63.7 | 1036.1 | <LOD | <LOD | 4.8 | 1293.7 | <LOD | 10.7 | <LOD | 718363
5 [Col] | 2023-07-31 | NIST 2711a | std | rep 5 | 14976.1 | 71722.3 | 268740.2 | <LOD | <LOD | 26303.7 | 24051.2 | 3229.9 | 164.5 | 114.5 | 576.3 | 28453 | <LOD | 28.1 | 176.1 | 396.1 | 121.6 | 1.7 | 116 | 252.7 | 32.6 | 284.3 | 17.7 | 6.2 | 27.7 | 80.4 | 66.2 | 88 | 620.7 | 19.1 | <LOD | 6.1 | 1353.4 | <LOD | 18.2 | <LOD | 720158.5
6 [Col] | 2023-07-31 | NIST 679 | std | rep 1 | 9769 | 89984.2 | 217554.2 | <LOD | <LOD | 23425.3 | 1519.1 | 5332.2 | 140.6 | <LOD | 1674.5 | 93830.3 | 188.1 | 36.8 | 83.7 | 121.7 | 23.5 | 5.4 | 181.8 | 45.3 | 38.8 | 163 | 24.5 | 8.9 | 25.1 | 30.9 | 69.8 | 28.2 | 720.3 | 52.3 | <LOD | <LOD | 4 | <LOD | <LOD | 4.2 | 708046.1
7 [Col] | 2023-07-31 | NIST 679 | std | rep 2 | 9769 | 88735.2 | 217199.4 | <LOD | <LOD | 24070.1 | 1565.6 | 5352.6 | <LOD | 106.8 | 1667.1 | 92355.9 | 265.2 | 43.6 | 60.8 | 104.6 | 17.2 | 2.1 | 186.6 | 52.2 | 45.8 | 168.9 | 22.7 | 20.9 | 53.2 | 57.3 | 71.9 | 118.5 | 1329.7 | 54.8 | <LOD | <LOD | 18.9 | <LOD | 27.8 | 6.7 | 709568.3
8 [Col] | 2023-07-31 | NIST 679 | std | rep 3 | 9769 | 89633.5 | 219551.7 | <LOD | <LOD | 23684 | 1457.3 | 4556 | 72.5 | 52.3 | 1651.1 | 92242.3 | 180.6 | 27.3 | 30.6 | 113.6 | 19.1 | 2.3 | 190 | 47.9 | 35.1 | 152.9 | 21.8 | 7 | 21.5 | <LOD | 58.7 | 142.5 | 1683.3 | 18.7 | <LOD | <LOD | 12.6 | <LOD | 29.3 | <LOD | 707894.7
9 [Col] | 2023-07-31 | NIST 679 | std | rep 4 | 9769 | 86618 | 216105.1 | <LOD | <LOD | 23349.8 | 1519 | 5142.6 | 107.5 | 80.4 | 1513.9 | 87108.9 | 375.3 | 28 | 78.5 | 115.8 | 23.1 | 3.2 | 180.1 | 45.1 | 38.7 | 146.1 | 13.3 | 12.6 | 32.5 | 71.6 | 72.3 | 59.4 | 1479.7 | <LOD | <LOD | <LOD | 9.7 | <LOD | 32.7 | <LOD | 718165.8
10 [Col] | 2023-07-31 | NIST 679 | std | rep 5 | 11982.6 | 88737.1 | 218700.1 | <LOD | <LOD | 23745.8 | 1447.8 | 6006.9 | <LOD | 80.3 | 1588.8 | 87990.1 | 260.1 | 53 | 48.2 | 125.1 | 11.4 | 4.4 | 182.8 | 53.3 | 40.1 | 150.6 | 22.3 | 20.6 | 31.5 | 20.1 | 72.8 | 65.5 | <LOD | <LOD | <LOD | <LOD | 15.1 | <LOD | 15 | <LOD | 711190.1
11 [Col] | 2023-07-31 | SARM 69 | std | rep 1 | 9769 | 79809.1 | 273396.7 | 174.4 | <LOD | 17892.1 | 14871.3 | 4675.6 | 285.1 | 174.8 | 951.5 | 51263 | 195.6 | 57.4 | 94.8 | 80.4 | 6.7 | 3.5 | 78.3 | 99 | 32.5 | 301.8 | 7.1 | 13.2 | 31.5 | 40.1 | 33.4 | 52.2 | 416.9 | <LOD | <LOD | <LOD | 12.6 | 25.5 | <LOD | <LOD | 703392
12 [Col] | 2023-07-31 | SARM 69 | std | rep 2 | 12466.3 | 79839.8 | 275590.9 | 214.6 | <LOD | 18162.1 | 15119.5 | 4471.6 | 156.1 | 85.4 | 1021.8 | 52636.5 | 92.3 | 55.5 | 68.1 | 69 | 7 | 4.5 | 73.9 | 99.7 | 29.8 | 307.8 | 6.7 | 15 | 25.5 | 48 | 45.4 | 69.8 | 940.2 | 18.9 | <LOD | <LOD | 15.1 | <LOD | 19.7 | 6 | 696404.9
13 [Col] | 2023-07-31 | SARM 69 | std | rep 3 | 11503.8 | 79102.5 | 273846.4 | 183.9 | <LOD | 17480.9 | 15017 | 4695 | 358.6 | 116.2 | 1012 | 49654 | 309.9 | 57.9 | 64.7 | 68.3 | 7 | 4 | 77.2 | 100.6 | 30.6 | 240.8 | 17 | 12.3 | <LOD | 28.7 | 34.1 | 51.1 | <LOD | <LOD | <LOD | <LOD | 14.9 | <LOD | 7.8 | <LOD | 703627.4
14 [Col] | 2023-07-31 | SARM 69 | std | rep 4 | 12680.5 | 77637.8 | 268016.8 | 200.7 | <LOD | 16973.2 | 14587 | 4004.3 | 231.3 | 253.7 | 918.7 | 48230.5 | 203.1 | 40 | 59.2 | 57.4 | 3.8 | 3.3 | 69.8 | 112.8 | 28.7 | 257.6 | 9.9 | 14.6 | 20.1 | 36.5 | 46.5 | 92.1 | 502.8 | 17.9 | <LOD | <LOD | 13.6 | <LOD | <LOD | <LOD | 711133.7
15 [Col] | 2023-07-31 | SARM 69 | std | rep 5 | 9769 | 52187 | 183229.4 | 51.2 | <LOD | 9859.6 | 6761.3 | 2140.5 | 41.2 | 59.4 | 409.9 | 22205.3 | <LOD | 10.1 | 26.9 | 22.7 | 1.2 | 2.6 | 30.6 | <LOD | 10.4 | 66.5 | 2.1 | 5.8 | <LOD | <LOD | <LOD | 23.6 | 483.5 | 6.6 | <LOD | <LOD | 4 | <LOD | 5 | <LOD | 851772.3
2 [Col] | 2023-08-01 | NIST 2711a | std | rep 1 | 11843.7 | 68700.7 | 264254.4 | <LOD | <LOD | 25573.2 | 23266.5 | 3299.8 | <LOD | 64.5 | 697.1 | 28108.3 | <LOD | 25.6 | 152 | 401.2 | 143.9 | <LOD | 117.8 | 251.6 | 31.4 | 332.6 | 14.5 | 4.1 | 16.3 | 81.8 | 78.5 | 82.3 | 1652.1 | <LOD | 6.9 | 13 | 1289 | 33.5 | <LOD | <LOD | 730576.2
3 [Col] | 2023-08-01 | NIST 2711a | std | rep 2 | 9769 | 70251.3 | 267427.2 | <LOD | <LOD | 25509.2 | 23212 | 2777.9 | 98.4 | 111.2 | 597.4 | 28949.6 | <LOD | 19.5 | 177.7 | 413.3 | 137.9 | 4.6 | 124.2 | 260.5 | 39.7 | 303.7 | 19.5 | 10.7 | 32.6 | 109.9 | <LOD | 29.2 | 1582.9 | <LOD | <LOD | <LOD | 1324.2 | <LOD | 10 | 6 | 728713.1
4 [Col] | 2023-08-01 | NIST 2711a | std | rep 3 | 13338.4 | 71830 | 264324.8 | <LOD | <LOD | 25853.1 | 23069.1 | 3686.3 | 163.8 | 122.5 | 669.6 | 28173.8 | 160.5 | 25.1 | 173.6 | 381.6 | 124.8 | <LOD | 114.7 | 257.7 | 30.5 | 279 | 15.3 | 6.8 | 20.8 | 69.2 | <LOD | 53.8 | 977.5 | <LOD | 5.7 | <LOD | 1303.2 | <LOD | <LOD | <LOD | 726343.8
5 [Col] | 2023-08-01 | NIST 2711a | std | rep 4 | 15708.5 | 70646.7 | 267313.9 | <LOD | <LOD | 26339.3 | 23682.4 | 3502.5 | 122.3 | 46 | 728 | 28834 | 89.2 | 22.6 | 179.5 | 413.6 | 136.6 | 5.4 | 117.7 | 261.6 | 44.5 | 322.4 | 23.6 | 7.7 | 33.7 | 91.2 | 44.3 | 47.4 | 1019.5 | <LOD | <LOD | <LOD | 1369.1 | 21.1 | <LOD | <LOD | 720410.7
6 [Col] | 2023-08-01 | NIST 2711a | std | rep 5 | 12120.7 | 71407.3 | 268031.3 | 64.2 | <LOD | 26422 | 24401.2 | 3912 | 150.1 | 58.2 | 679.7 | 29438 | 101.4 | 25.9 | 160.6 | 422.5 | 147.8 | <LOD | 122.8 | 267.4 | 34.8 | 334.3 | 21.2 | 5.5 | <LOD | 43.8 | 47.5 | 46.2 | 390.8 | 42.5 | <LOD | 8.4 | 1353.2 | <LOD | 18.3 | <LOD | 722389.5
7 [Col] | 2023-08-01 | NIST 679 | std | rep 1 | 9769 | 93168.4 | 221370.9 | <LOD | <LOD | 24086.2 | 1523.9 | 5328.1 | 145.2 | 62.8 | 1523.7 | 92412.5 | 172.9 | 56.9 | 57.3 | 144.7 | 21.1 | 4.5 | 187.4 | 50.6 | 36 | 164.2 | 14 | 10.4 | 24.9 | <LOD | 80 | 62.1 | 1011.4 | <LOD | <LOD | 8.1 | 12.6 | <LOD | 28.9 | <LOD | 702858.2
8 [Col] | 2023-08-01 | NIST 679 | std | rep 2 | 9769 | 91607.8 | 219989.6 | <LOD | <LOD | 23919.2 | 1474.7 | 5399.4 | 90.7 | 84.8 | 1737.3 | 91561.9 | <LOD | 43.3 | 32 | 123.1 | 27.5 | 3.3 | 182 | 42.3 | 35.9 | 154.1 | 14.6 | 15.4 | <LOD | 25.1 | 40.4 | 54.8 | 1154 | 28 | <LOD | <LOD | 11.4 | <LOD | 14.6 | <LOD | 706225.1
9 [Col] | 2023-08-01 | NIST 679 | std | rep 3 | 9769 | 94451.2 | 225959 | <LOD | <LOD | 24593.6 | 1587.4 | 5322.5 | 385 | 62.1 | 1707.1 | 91478.7 | 139.6 | 61.1 | 49.1 | 122.4 | 30.1 | 3.4 | 196.5 | 48.2 | 40 | 160.7 | 14.3 | <LOD | 37.7 | 38.6 | 73.4 | 55.7 | 348.3 | 30.3 | <LOD | <LOD | 8.7 | <LOD | 30.4 | <LOD | 698463.5
10 [Col] | 2023-08-01 | NIST 679 | std | rep 4 | 9769 | 92202.6 | 222986.3 | <LOD | <LOD | 24206.9 | 1548.3 | 5005.3 | 197.9 | 67.6 | 1615.2 | 92459.4 | 243.7 | 62.9 | 83.4 | 120.5 | 19.7 | 7.1 | 186.5 | 52.5 | 32.8 | 150 | 21.7 | 18.3 | 26.1 | 45.6 | 53.2 | 55.1 | 1731.9 | <LOD | <LOD | 9.6 | 13.7 | <LOD | 20.7 | <LOD | 701383.5
11 [Col] | 2023-08-01 | NIST 679 | std | rep 5 | 12465.8 | 88922.6 | 216343.9 | <LOD | <LOD | 23753.1 | 1551.3 | 5339.9 | 130.9 | 75.2 | 1556.9 | 89229.1 | 179.8 | 53.4 | 62.3 | 114.6 | 16.8 | 5.9 | 189.4 | 46.9 | 42.3 | 158.4 | 8.1 | <LOD | <LOD | 16.8 | 66.5 | 105.7 | 1372.8 | 35.2 | <LOD | 7.8 | 16.2 | <LOD | 35.2 | <LOD | 710621.7
12 [Col] | 2023-08-01 | SARM 69 | std | rep 1 | 11314 | 75952.4 | 262828.5 | 134.6 | <LOD | 17176.5 | 14046.4 | 4530.5 | 103.7 | 161.2 | 949.6 | 51165.8 | 244.2 | 50.1 | 61.3 | 47.1 | 6.2 | 3.5 | 74.1 | 108.6 | 35.8 | 329.7 | 11.8 | 6.7 | 14.5 | 20.9 | 40.2 | 58.8 | 1787.5 | 56.4 | <LOD | <LOD | 16.8 | 35.2 | <LOD | <LOD | 714367.8
13 [Col] | 2023-08-01 | SARM 69 | std | rep 2 | 13411.2 | 79157.7 | 268087.8 | 83.9 | <LOD | 17563.2 | 14462.4 | 5371.1 | 93.1 | 198.5 | 993.5 | 52934.1 | 109.3 | 49 | 73.3 | 83.7 | 5.7 | 1.9 | 78.4 | 114.8 | 35.2 | 313.1 | 3.7 | 5.9 | <LOD | 40.2 | 61.1 | 61.8 | 930.1 | 19.8 | <LOD | <LOD | 15.9 | 25.2 | <LOD | <LOD | 702356.9
14 [Col] | 2023-08-01 | SARM 69 | std | rep 3 | 12100.2 | 78886.3 | 271695.6 | 158.2 | <LOD | 17329 | 14509 | 4670.8 | 181.8 | 108.1 | 1033.6 | 51405.9 | 79.9 | 59.6 | 61.2 | 56.4 | 10.5 | 5.6 | 81.3 | 112.6 | 34.4 | 268.2 | 10.1 | 17.4 | 12.6 | 44 | 46.3 | 47.5 | 712.5 | 24.1 | <LOD | <LOD | 13.1 | 52.1 | <LOD | 5.4 | 703324.5
15 [Col] | 2023-08-01 | SARM 69 | std | rep 4 | 11440.9 | 77238.6 | 270726.4 | 116.4 | <LOD | 17827.7 | 14660.7 | 4715.9 | 117.1 | 251.7 | 925.6 | 51576.4 | <LOD | 67.3 | 84.5 | 63.1 | 7.5 | 3.6 | 79.2 | 97.4 | 31.1 | 271.4 | 12.9 | 11.6 | <LOD | 28.8 | 52.8 | <LOD | 664.9 | 40.7 | <LOD | <LOD | 14.4 | <LOD | 7.8 | <LOD | 705949.6
16 [Col] | 2023-08-01 | SARM 69 | std | rep 5 | 13479.1 | 79346.6 | 271069.4 | 217.7 | <LOD | 17895.5 | 14634.7 | 4181.8 | 194.9 | 164.7 | 885.3 | 51153.6 | 62.6 | 32.4 | 58.7 | 58.9 | 10.6 | 3.3 | 77.9 | 93.3 | 35.1 | 259 | 9.3 | 16.3 | <LOD | 43 | 32.5 | 42.6 | 1277.1 | 32.2 | <LOD | 6.8 | 10.9 | <LOD | 9.7 | <LOD | 701909.2
2 [Col] | 2023-08-02 | NIST 2711a | std | rep 1 | 14904.3 | 72373.6 | 267316.1 | 63.9 | <LOD | 25672.6 | 23292.2 | 3755.4 | 160.3 | <LOD | 721.4 | 29710.7 | 93.7 | 21.5 | 195.3 | 408.8 | 125.7 | 2.4 | 117.5 | 272.3 | 31.4 | 321.6 | 19.3 | 11 | 35.7 | 84.2 | 75.9 | 106.4 | 633.6 | <LOD | <LOD | 13.7 | 1324.3 | <LOD | 10.7 | <LOD | 719786.7
3 [Col] | 2023-08-02 | NIST 2711a | std | rep 2 | 13846 | 68572 | 265496.3 | <LOD | <LOD | 25726.5 | 22985.6 | 3413.4 | 128.7 | <LOD | 562.1 | 28755 | 75.5 | 11.4 | 158.1 | 404.2 | 129.8 | 4.7 | 123.2 | 266.5 | 32.8 | 303.3 | 21.8 | 8.7 | 15.6 | 96.9 | 38.3 | 62.9 | 1701.8 | <LOD | 5.9 | <LOD | 1295.3 | 26.1 | <LOD | <LOD | 726537.3
4 [Col] | 2023-08-02 | NIST 2711a | std | rep 3 | 15486.9 | 69826.7 | 265869.5 | <LOD | <LOD | 25236.5 | 23313.5 | 4110.9 | 133.4 | 64.2 | 642.3 | 28957.6 | 127.5 | 16.9 | 167 | 394.5 | 93.4 | 3.3 | 109.7 | 270.9 | 38.6 | 299.7 | 22.4 | <LOD | 16.2 | 90.8 | 28.5 | 88.5 | 838 | <LOD | 6.2 | 7.8 | 1301.8 | <LOD | <LOD | <LOD | 723147.2
5 [Col] | 2023-08-02 | NIST 2711a | std | rep 4 | 13166 | 70375.6 | 273831.8 | <LOD | <LOD | 26689.4 | 23746.1 | 3861.6 | <LOD | <LOD | 657.4 | 30413.2 | 103.3 | 22.7 | 177.2 | 397.3 | 169.2 | 3.6 | 117.3 | 261.7 | 30.3 | 288.8 | 18.4 | 15.8 | <LOD | 93.2 | <LOD | 80.6 | 2215.8 | 44.7 | <LOD | 15.2 | 1365.5 | <LOD | 20.5 | 8.1 | 714400.6
6 [Col] | 2023-08-02 | NIST 2711a | std | rep 5 | 13497.5 | 70781.5 | 270252.5 | <LOD | <LOD | 26623.9 | 23924.9 | 3630.2 | 204.6 | 34.7 | 606.2 | 29020.2 | <LOD | 15.1 | 154.6 | 397.8 | 127.9 | 6.4 | 116.5 | 261.8 | 34.4 | 284.7 | 23 | 6.6 | 27.8 | 104.7 | 38.8 | 56 | 1025.7 | 44.4 | <LOD | <LOD | 1320 | <LOD | <LOD | <LOD | 719794.5
7 [Col] | 2023-08-02 | NIST 679 | std | rep 1 | 9769 | 91872.5 | 223046.5 | <LOD | <LOD | 24141 | 1455 | 5173.9 | 197.3 | 158.7 | 1442.5 | 89659.9 | 373.7 | 60.8 | 74.5 | 124.2 | 30.5 | 2.9 | 189.4 | 52.1 | 42.5 | 160.1 | 8.4 | 7 | <LOD | 34.7 | 97.5 | 130.9 | 1026.5 | 26.7 | <LOD | <LOD | 12.2 | <LOD | 27.4 | <LOD | 704868.4
8 [Col] | 2023-08-02 | NIST 679 | std | rep 2 | 9769 | 90330.7 | 218594.9 | <LOD | <LOD | 24268.3 | 1559.9 | 5656.4 | 84.7 | 40.8 | 1627.5 | 93097.5 | 96 | 78.7 | 39.2 | 119.3 | 14 | 5.7 | 189.9 | 45.1 | 33.8 | 157.7 | 23.6 | 10.4 | 23.2 | 53.4 | 55.4 | 108.5 | 913.4 | 20.4 | <LOD | <LOD | 8.8 | 20.5 | <LOD | <LOD | 706575.4
9 [Col] | 2023-08-02 | NIST 679 | std | rep 3 | 9769 | 94513.5 | 227128 | <LOD | <LOD | 24955.6 | 1642.4 | 5744.4 | 117.8 | 98.4 | 1714.4 | 93163.2 | 211 | 50.2 | 104.9 | 125.7 | 18.8 | 2.7 | 197.6 | 49.4 | 36.2 | 167.9 | 27.5 | 13.7 | 26 | 31.7 | 54.2 | 56.5 | 492.6 | 52.1 | <LOD | <LOD | 16.4 | <LOD | 13.5 | <LOD | 694911.3
10 [Col] | 2023-08-02 | NIST 679 | std | rep 4 | 9769 | 91120 | 221946.5 | <LOD | <LOD | 24655 | 1482.5 | 5645.7 | 86.9 | 70.4 | 1648.4 | 91807.2 | <LOD | 23.3 | 80.2 | 137 | 15.6 | 5.5 | 184.5 | 60.7 | 27.2 | 162.8 | 23.8 | 14.2 | 25.5 | 43.6 | 125.4 | 111.4 | <LOD | <LOD | <LOD | 5.5 | 23.5 | <LOD | 35.7 | <LOD | 704864.4
11 [Col] | 2023-08-02 | NIST 679 | std | rep 5 | 9769 | 96929.8 | 227238.5 | <LOD | <LOD | 24690.7 | 1465.6 | 5189.7 | 82 | 52.5 | 1725.4 | 91240.5 | 284.7 | 45.8 | 48.6 | 127.8 | 23.2 | 7.9 | 191.4 | 53.7 | 48.2 | 157.7 | 21.5 | 11.8 | <LOD | 37.3 | 54.5 | 127.1 | 1163.9 | 20.2 | <LOD | <LOD | 13.5 | <LOD | 24 | <LOD | 695069.7
12 [Col] | 2023-08-02 | SARM 69 | std | rep 1 | 12911.4 | 80094.2 | 275286.3 | 162.3 | <LOD | 17828.4 | 15269.3 | 4763.9 | 210.4 | 209.2 | 969.9 | 51632.9 | 106.3 | 59.4 | 79.2 | 67.8 | 4.5 | 4.4 | 74.5 | 96 | 35.5 | 283.4 | 5.8 | 10.3 | <LOD | <LOD | <LOD | 66 | 383.7 | 43.1 | <LOD | <LOD | 19.8 | <LOD | <LOD | <LOD | 697367.6
13 [Col] | 2023-08-02 | SARM 69 | std | rep 2 | 11314 | 75451.1 | 267164.3 | 156.5 | <LOD | 17252.8 | 14497.8 | 4491.1 | 277.5 | 119.9 | 993.1 | 49725.5 | 130.5 | 50.6 | 80.2 | 74.3 | 10.3 | 4 | 70.2 | 106 | 31.2 | 306.8 | 11.7 | 12.3 | 12.7 | 23.9 | 53.2 | 25.6 | 790.4 | <LOD | <LOD | 4.9 | 8.6 | 28.7 | <LOD | <LOD | 712889.9
14 [Col] | 2023-08-02 | SARM 69 | std | rep 3 | 12876.9 | 79211.8 | 273694.6 | 187.8 | <LOD | 17662.2 | 15221.4 | 5024.1 | 193 | 168.9 | 891.9 | 51612.6 | 314.7 | 81.4 | 64.8 | 66.1 | 10.9 | 3.5 | 78.3 | 104.9 | 31.9 | 291.9 | 11.2 | <LOD | 42.6 | 75.6 | 68.5 | 49.3 | 757.9 | <LOD | <LOD | <LOD | 9.2 | 35.8 | <LOD | 3.8 | 698783.6
15 [Col] | 2023-08-02 | SARM 69 | std | rep 4 | 12073.2 | 80963.3 | 275159.3 | 230.7 | <LOD | 17684 | 15415.2 | 4741.6 | 115.1 | 187.1 | 958.7 | 51234.6 | <LOD | 71 | 83.8 | 67.9 | 9 | 2.1 | 73.8 | 104.9 | 33.2 | 267.3 | 6.8 | 12.2 | 26 | 39 | 49.8 | 76.8 | 931.6 | 13.5 | <LOD | <LOD | 15.8 | <LOD | <LOD | <LOD | 697715.3
16 [Col] | 2023-08-02 | SARM 69 | std | rep 5 | 12482.3 | 79668.9 | 272626.5 | 176.6 | <LOD | 17794 | 14868.7 | 4573.6 | 105.4 | 142.6 | 994.2 | 50938 | 342.4 | 58.5 | 77.9 | 64.5 | 4.8 | 3.3 | 74.4 | 98.6 | 29.2 | 262 | 9.1 | 9.6 | 12.5 | <LOD | 54.2 | 66.6 | 1154.2 | 25.6 | <LOD | <LOD | 15.7 | 20.1 | <LOD | <LOD | 700917.4
2 [Col] | 2023-08-03 | NIST 2711a | std | rep 1 | 13256 | 72252.2 | 271126.3 | <LOD | <LOD | 26104.5 | 24448.5 | 2996.3 | 197.2 | <LOD | 787.5 | 29307 | 208.9 | 28.5 | 170.9 | 389.4 | 136 | <LOD | 119.4 | 258.7 | 34.7 | 251.9 | 17.4 | 14 | <LOD | 96.9 | 43.4 | 103.7 | 1212.6 | 40.5 | <LOD | <LOD | 1350.2 | <LOD | <LOD | <LOD | 717910.6
3 [Col] | 2023-08-03 | NIST 2711a | std | rep 2 | 13390.4 | 72898 | 272749.4 | <LOD | <LOD | 26431.3 | 24295.8 | 3071.2 | 132.7 | 44.2 | 696.1 | 28402 | 56.3 | 42.9 | 217.2 | 370.1 | 135.5 | <LOD | 118.5 | 263.7 | 41.4 | 353.5 | 13.8 | <LOD | <LOD | 66.3 | 47.9 | 32.8 | 426.6 | <LOD | <LOD | 15.5 | 1310.2 | 39.1 | <LOD | 4.8 | 717526.2
4 [Col] | 2023-08-03 | NIST 2711a | std | rep 3 | 15005.5 | 71096.6 | 271226.3 | <LOD | <LOD | 26485.3 | 24377.8 | 3403.7 | 71.2 | 35.9 | 673.6 | 29044.8 | 46.7 | 27.2 | 208.8 | 381.1 | 138.2 | 1.8 | 115.5 | 275.4 | 33.5 | 341.4 | 16.2 | 14.1 | <LOD | 40.6 | 19.4 | 70.9 | 779.6 | 17.3 | 5.4 | 7.9 | 1385.8 | <LOD | 8.7 | <LOD | 717084.7
5 [Col] | 2023-08-03 | NIST 2711a | std | rep 4 | 13690.3 | 71131.2 | 269403.4 | <LOD | <LOD | 26384.8 | 24049.9 | 3229.5 | 119.7 | 70.8 | 735.7 | 30386.7 | 82.3 | 31.9 | 177.6 | 431.2 | 156 | 3 | 116.5 | 270.1 | 38.6 | 289.2 | 24.9 | 20.5 | 16.2 | 58.2 | 27.6 | 115.4 | 900.7 | <LOD | <LOD | 10.6 | 1331.2 | <LOD | <LOD | <LOD | 719050.8
6 [Col] | 2023-08-03 | NIST 2711a | std | rep 5 | 13943.6 | 71189.6 | 271424.7 | <LOD | <LOD | 26253.2 | 23737.1 | 3500.7 | 92.1 | 45.8 | 588.2 | 28682.2 | 128.6 | 12.5 | 180.9 | 396.1 | 136.4 | 7 | 123.4 | 261 | 33.5 | 268.3 | 16.5 | 9 | 14 | 72.4 | 18.1 | 40.4 | 720.9 | <LOD | <LOD | 7 | 1331.8 | 17.1 | 16.5 | <LOD | 719040.5
7 [Col] | 2023-08-03 | NIST 679 | std | rep 1 | 9769 | 94568 | 225547 | <LOD | <LOD | 24619 | 1540.2 | 5839.9 | 320 | 76.3 | 1388.4 | 91467.4 | 217.4 | 38.1 | 62.6 | 125.1 | 15.4 | 5.1 | 190.9 | 62.3 | 48.4 | 158.1 | 10.7 | 4.2 | 27.6 | <LOD | 33.8 | 75.3 | <LOD | <LOD | <LOD | <LOD | 13.7 | <LOD | 31.6 | 6.9 | 698966.9
8 [Col] | 2023-08-03 | NIST 679 | std | rep 2 | 9769 | 91842 | 220225.3 | <LOD | <LOD | 24001.8 | 1626.8 | 5565.9 | 201 | 126.4 | 1704.1 | 92040.4 | <LOD | 51.9 | 80.7 | 127.4 | 16.9 | 1.8 | 187.7 | 49 | 42.2 | 174.1 | 22.1 | 10.4 | 41.8 | 62.2 | 37.5 | 69.6 | 659.5 | 27 | <LOD | <LOD | 8.1 | <LOD | <LOD | <LOD | 705234.5
9 [Col] | 2023-08-03 | NIST 679 | std | rep 3 | 9769 | 94413.3 | 226600.7 | <LOD | <LOD | 24768 | 1515.6 | 5338.8 | 244.9 | 71.5 | 1692.8 | 95273.2 | <LOD | 68.8 | 49.2 | 125.7 | 30.1 | 2.6 | 179.8 | 48.1 | 32.9 | 150.7 | 19.6 | 10 | 18.7 | 51.6 | 62.7 | 110.2 | 794.2 | <LOD | <LOD | <LOD | 4 | <LOD | 53.3 | <LOD | 693840.3
10 [Col] | 2023-08-03 | NIST 679 | std | rep 4 | 9769 | 96429.1 | 228069 | <LOD | <LOD | 25379 | 1679.9 | 5958.9 | 187.6 | 101.6 | 1857.6 | 93325 | <LOD | 35.2 | 87.8 | 123.5 | 24.9 | 3.6 | 201.9 | 48.8 | 38.3 | 165.7 | 18.6 | 13.1 | 37.9 | 18 | 55.1 | 92.3 | <LOD | 55.7 | <LOD | <LOD | 11.9 | <LOD | 38.1 | <LOD | 692341.8
11 [Col] | 2023-08-03 | NIST 679 | std | rep 5 | 9769 | 92394.1 | 224719.3 | <LOD | <LOD | 24825.5 | 1563.3 | 5762.1 | 162.7 | 56.6 | 1507.3 | 93611.7 | 174.2 | 39.1 | 60 | 125.5 | 21.4 | 7 | 200.6 | 51.5 | 39.9 | 152.8 | 14 | 10.1 | <LOD | 22.4 | 43.1 | 65.7 | <LOD | 44.3 | <LOD | <LOD | 15.6 | <LOD | 40.1 | 5 | 699270.6
12 [Col] | 2023-08-03 | SARM 69 | std | rep 1 | 13206.4 | 79680.3 | 273349.9 | 160.9 | <LOD | 17850.9 | 14710.9 | 4464.2 | 315.4 | 258.6 | 965.6 | 51543.6 | 290.6 | 59.6 | 69.4 | 72.6 | 6 | 3.3 | 76.7 | 105.6 | 30.8 | 292.9 | 13.8 | 15 | 13.6 | 39.5 | 67.4 | 82.7 | 1199.9 | 16.4 | <LOD | <LOD | 17 | 33 | <LOD | 7 | 698561.9
13 [Col] | 2023-08-03 | SARM 69 | std | rep 2 | 12033.3 | 80845.2 | 277054.4 | 193 | <LOD | 18043.1 | 15100.6 | 5079.4 | 138.1 | 289.9 | 991.6 | 52495.4 | 88.4 | 66.9 | 95.2 | 88.6 | 14 | 3.9 | 79.7 | 96.7 | 33.1 | 341.3 | 6.6 | 5.1 | 24.1 | 37.9 | 57.5 | 131.6 | 439.4 | <LOD | <LOD | 8.3 | 12.6 | <LOD | 23.9 | <LOD | 694565.7
14 [Col] | 2023-08-03 | SARM 69 | std | rep 3 | 12111.3 | 79589.8 | 266934 | 168.5 | <LOD | 17500 | 14685.7 | 5180.9 | 101.5 | 271.2 | 971.3 | 52425.6 | 137.4 | 63 | 81 | 77 | 9.3 | 5.9 | 78 | 96.6 | 25.4 | 297.1 | 14.5 | 9.2 | 17.7 | 52.6 | 58.7 | 64.4 | <LOD | 16.6 | <LOD | 6.3 | 12.7 | <LOD | 10.1 | <LOD | 705958.6
15 [Col] | 2023-08-03 | SARM 69 | std | rep 4 | 13215.8 | 79105 | 270687 | 173.7 | <LOD | 17935.4 | 15101.3 | 4497.5 | <LOD | 142.7 | 818.6 | 49860.8 | 82.7 | 58.5 | 64.1 | 74.4 | <LOD | 4.3 | 71 | 90 | 26.7 | 306 | 8.6 | 7.8 | 14.3 | 15.4 | 43.2 | 104.8 | 1264.1 | 14.1 | <LOD | <LOD | 19.4 | <LOD | <LOD | 3.3 | 703623.8
16 [Col] | 2023-08-03 | SARM 69 | std | rep 5 | 12103.6 | 80910.9 | 277523.9 | 192.2 | <LOD | 17877.5 | 14888.3 | 4590.6 | 123.8 | 273.1 | 884.8 | 50384.8 | <LOD | 58.4 | 69.5 | 63.9 | 7.1 | <LOD | 74.4 | 104.8 | 28 | 274.2 | 7.9 | 8.4 | <LOD | 26.9 | 39.6 | 37.4 | 423.9 | 23.9 | <LOD | <LOD | 17.9 | 22.8 | <LOD | <LOD | 697389.2
2 [Col] | 2023-08-22 | NIST2711a | standard | rep 1 | 14759.7 | 70232 | 268660.8 | <LOD | <LOD | 26393 | 23816.1 | 3677.4 | 82.9 | <LOD | 683.3 | 29995.6 | 132.2 | 14.1 | 168.3 | 432.4 | 127.5 | <LOD | 121.1 | 277.9 | 36 | 272.6 | 25.5 | 18.8 | 21.2 | 60.5 | 63.2 | 106.8 | 1269.2 | 17.7 | <LOD | 13.3 | 1394.8 | <LOD | 25 | <LOD | 718910
3 [Col] | 2023-08-22 | NIST2711a | standard | rep 2 | 14426.3 | 69389.7 | 265666.6 | <LOD | <LOD | 25814.6 | 23732.2 | 3057.1 | <LOD | 59.7 | 630.3 | 27818.7 | 165.1 | 19.8 | 167.6 | 402.4 | 115.1 | 5.2 | 120 | 250.9 | 35.6 | 302.4 | 21 | 7.1 | 17.9 | 52.7 | 30.3 | 68.1 | 523.3 | <LOD | <LOD | 13.2 | 1282.4 | <LOD | 12 | <LOD | 727036.1
4 [Col] | 2023-08-22 | NIST2711a | standard | rep 3 | 12754.6 | 71234.9 | 270875.5 | <LOD | <LOD | 26686.1 | 24899.1 | 3341.5 | <LOD | <LOD | 643.2 | 31171 | 107.5 | 13.4 | 186.3 | 414.7 | 164.4 | <LOD | 125 | 286.2 | 35.9 | 258.1 | 17 | 7.3 | 15 | 91 | 64.2 | 87.5 | 1554.2 | 39.5 | <LOD | 13.3 | 1383.2 | <LOD | 26.7 | <LOD | 716585.5
5 [Col] | 2023-08-22 | NIST2711a | standard | rep 4 | 13523.7 | 72131.1 | 273188.4 | <LOD | <LOD | 26731.6 | 25128.7 | 3236.9 | 174.4 | 37.2 | 563.6 | 30852 | <LOD | 12.8 | 165.8 | 422.9 | 129.4 | 2 | 121 | 263.7 | 27.6 | 322.4 | 24.7 | 13 | 42.3 | 81.5 | 76.7 | 96.7 | 412.5 | 17.4 | <LOD | 15.6 | 1393.1 | <LOD | 7.5 | <LOD | 714227.1
6 [Col] | 2023-08-22 | NIST2711a | standard | rep 5 | 15314.6 | 69374.1 | 268264.5 | <LOD | <LOD | 26715.8 | 23750.7 | 3051.1 | 105.1 | 47.5 | 668 | 28665.7 | <LOD | 28.7 | 178 | 376.6 | 145.4 | 3.2 | 117.2 | 270.2 | 34 | 292.9 | 22.8 | 8.9 | 59.7 | 74.9 | 33.5 | 51.8 | 904.5 | 53.5 | <LOD | <LOD | 1360.8 | <LOD | 8.6 | 8 | 721682
7 [Col] | 2023-08-22 | NIST 679 | standard | rep 1 | 9769 | 91548.2 | 220027.9 | <LOD | <LOD | 24321 | 1647.3 | 5430.7 | 228.1 | 92 | 1706.9 | 92568.5 | <LOD | 61.2 | 65.8 | 116.2 | 22.5 | 10.2 | 188 | 53.6 | 40.2 | 152.1 | 19.7 | 14.3 | <LOD | 35.4 | 38.3 | 79.4 | 925.4 | 37 | <LOD | 8.3 | 15.9 | <LOD | 8.4 | 4.2 | 704838.9
8 [Col] | 2023-08-22 | NIST 679 | standard | rep 2 | 9769 | 90270.6 | 220672 | <LOD | <LOD | 24199.8 | 1641.3 | 5421.8 | 100.1 | 34.2 | 1702.4 | 92502.9 | 168.3 | 80 | 67 | 123.3 | 27.2 | 7.5 | 185.7 | 54.6 | 33.1 | 161.7 | 25.3 | 16.9 | 49.8 | 51.1 | 85.9 | 111.8 | 757.9 | 54.3 | <LOD | <LOD | 13.7 | 22 | <LOD | <LOD | 705380.6
9 [Col] | 2023-08-22 | NIST 679 | standard | rep 3 | 9769 | 89989.8 | 220593.2 | <LOD | <LOD | 24447 | 1510.7 | 5801.8 | 190.5 | 32.1 | 1630.9 | 93159.4 | 97 | 70.2 | 90.6 | 122.1 | 23.3 | 6 | 202.1 | 49.5 | 43.4 | 173.6 | 22.9 | 9.7 | 27.7 | 36.8 | 52.5 | 110.9 | <LOD | <LOD | <LOD | <LOD | 20.2 | <LOD | 37.8 | <LOD | 705423.2
10 [Col] | 2023-08-22 | NIST 679 | standard | rep 4 | 9769 | 92589.1 | 221543.6 | <LOD | <LOD | 23967 | 1395.2 | 5071.1 | 234.3 | <LOD | 1792.1 | 90363 | 240.2 | 48.4 | 56.3 | 127.5 | 26.2 | 4.4 | 187.8 | 50.3 | 36.4 | 157.4 | 20.2 | 6.3 | <LOD | <LOD | 44.6 | 47.3 | 847.6 | 20.7 | <LOD | 12.5 | 17.5 | <LOD | 24.1 | <LOD | 705512.5
11 [Col] | 2023-08-22 | NIST 679 | standard | rep 5 | 9769 | 86753.5 | 214972.9 | <LOD | <LOD | 23411.9 | 1555.3 | 5461.1 | 280.1 | 87.4 | 1534.2 | 90771.8 | 264.1 | 61.9 | 55.3 | 114.6 | 16.3 | 3.2 | 188.1 | 53.7 | 37.4 | 166.4 | 19.7 | 9.3 | <LOD | 23.6 | 43.6 | 105.5 | <LOD | <LOD | <LOD | 5.3 | 21.9 | <LOD | 9.1 | 5.4 | 716433.3
12 [Col] | 2023-08-22 | SARM 69 | standard | rep 1 | 12959.4 | 78232.8 | 268080.7 | 152.8 | <LOD | 17555.2 | 14857.7 | 4707.2 | 126.4 | 152.1 | 1050.6 | 51302.7 | 173.6 | 71 | 54.6 | 62 | 7 | 1.5 | 78.1 | 104.8 | 28.6 | 251.1 | 9.7 | 9.5 | <LOD | 33.6 | 43.2 | 44.4 | 999.8 | <LOD | 7.6 | 9 | 21.9 | <LOD | <LOD | <LOD | 705621.3
13 [Col] | 2023-08-22 | SARM 69 | standard | rep 2 | 12577.8 | 77957.6 | 268298.5 | 151.8 | <LOD | 17717 | 14568.4 | 4505.2 | 215.5 | 262.1 | 963.2 | 50566.8 | 235 | 59.1 | 82.1 | 71.8 | 2.8 | 3 | 73.6 | 103.8 | 26.4 | 284.6 | 5.7 | <LOD | 49.7 | 45.8 | 30.3 | 56.9 | 1201.9 | <LOD | 3.8 | <LOD | 19.1 | <LOD | 7.3 | <LOD | 706662.1
14 [Col] | 2023-08-22 | SARM 69 | standard | rep 3 | 9769 | 79959.3 | 272471.9 | 188.9 | <LOD | 18405.7 | 14985.1 | 4337.9 | 327.4 | 177.1 | 940.2 | 52366.2 | <LOD | 66.9 | 102 | 71.7 | 10.1 | 2.9 | 76.7 | 97 | 27.1 | 307 | 5.1 | 10.1 | 18.9 | 20 | 36.2 | 54.6 | 635.2 | 15.7 | <LOD | <LOD | 17.6 | <LOD | 18.1 | 7.5 | 702944.3
15 [Col] | 2023-08-22 | SARM 69 | standard | rep 4 | 11559.5 | 78728.9 | 270051 | 200.3 | <LOD | 17752.4 | 15285.8 | 4774 | 343.3 | 103.2 | 904.4 | 51802.8 | 241.7 | 39.5 | 84.6 | 70.9 | 12.7 | 3.7 | 75.3 | 101.8 | 40.4 | 263.5 | 3.7 | 10.2 | 17.8 | 32.4 | 55.5 | 91.3 | 429.2 | <LOD | <LOD | <LOD | 13.9 | 16.8 | <LOD | <LOD | 704430.7
16 [Col] | 2023-08-22 | SARM 69 | standard | rep 5 | 13470.7 | 79686.8 | 275295.4 | 235.1 | <LOD | 18389.5 | 15353.8 | 5409.1 | <LOD | 194.7 | 1005.6 | 52765.9 | <LOD | 44.1 | 137.8 | 73.2 | 10.4 | <LOD | 76.6 | 105.6 | 33.4 | 306.2 | 11.5 | 11.4 | <LOD | 50.8 | 63.4 | 65.8 | 905.6 | <LOD | <LOD | <LOD | 16.3 | <LOD | 27.7 | <LOD | 694334.4
1 [Col] | 2024-05-14 | NIST 2711a | std | rep 1 | 14782 | 69584 | 265770 | <LOD | <LOD | 26958 | 24551 | 3181 | <LOD | 87 | 655 | 29057 | 50 | 17 | 163 | 363 | 132 | 5 | 119 | 251 | 31 | 351 | 25 | 7 | 12 | 80 | 23 | 82 | 1518 | 38 | <LOD | <LOD | 1319 | <LOD | <LOD | <LOD | 722799
2 [Col] | 2024-05-14 | NIST 2711a | std | rep 2 | 13594 | 69569 | 260279 | <LOD | <LOD | 26870 | 24961 | 2981 | 71 | 104 | 690 | 28836 | 92 | 21 | 178 | 398 | 153 | 3 | 124 | 252 | 28 | 303 | 18 | 10 | 17 | 67 | 39 | 87 | 1535 | 41 | <LOD | 10 | 1352 | <LOD | <LOD | <LOD | 729210
3 [Col] | 2024-05-14 | NIST 2711a | std | rep 3 | 12937 | 69986 | 266201 | <LOD | <LOD | 26372 | 24421 | 2891 | 126 | 97 | 613 | 27918 | 165 | 22 | 145 | 400 | 130 | 3 | 120 | 250 | 29 | 318 | 26 | 11 | 19 | 66 | 27 | 82 | 1962 | 30 | <LOD | 10 | 1294 | 36 | <LOD | <LOD | 725456
4 [Col] | 2024-05-14 | NIST 2711a | std | rep 4 | 14818 | 69948 | 267316 | <LOD | <LOD | 27258 | 25880 | 3342 | <LOD | 99 | 790 | 29792 | <LOD | 33 | 177 | 418 | 143 | <LOD | 121 | 267 | 29 | 320 | 14 | 5 | 16 | 87 | 86 | 95 | 1510 | <LOD | <LOD | 9 | 1379 | <LOD | 9 | 5 | 718820
5 [Col] | 2024-05-14 | NIST 2711a | std | rep 5 | 13830 | 70391 | 268520 | <LOD | <LOD | 26742 | 25026 | 3436 | 203 | <LOD | 608 | 29764 | <LOD | 38 | 177 | 388 | 123 | <LOD | 120 | 257 | 26 | 272 | 23 | 13 | <LOD | 79 | 31 | 80 | 1418 | <LOD | <LOD | <LOD | 1325 | <LOD | 7 | <LOD | 719716
6 [Col] | 2024-05-14 | NIST 679 | std | rep 1 | 9769 | 89060 | 218471 | 466 | <LOD | 24860 | 1920 | 5720 | 417 | 88 | 1537 | 88368 | 178 | 51 | 89 | 134 | 25 | 6 | 186 | 49 | 37 | 161 | 15 | <LOD | 15 | 64 | 131 | 80 | <LOD | 30 | 4 | <LOD | 4 | <LOD | 19 | <LOD | 711778
7 [Col] | 2024-05-14 | NIST 679 | std | rep 2 | 9769 | 88828 | 216347 | 462 | <LOD | 24403 | 1925 | 5176 | 174 | 84 | 1536 | 85986 | 270 | 47 | 62 | 123 | 18 | 4 | 189 | 44 | 35 | 155 | 16 | 11 | 42 | 32 | 43 | 62 | 548 | <LOD | <LOD | <LOD | 13 | 24 | <LOD | <LOD | 716953
8 [Col] | 2024-05-14 | NIST 679 | std | rep 3 | 9769 | 87222 | 214286 | 430 | <LOD | 24395 | 1930 | 5333 | 244 | 95 | 1658 | 87508 | 225 | 41 | 63 | 124 | 12 | <LOD | 190 | 48 | 40 | 165 | 18 | 10 | <LOD | <LOD | 72 | 74 | <LOD | <LOD | <LOD | <LOD | 21 | <LOD | 17 | <LOD | 718860
9 [Col] | 2024-05-14 | NIST 679 | std | rep 4 | 9769 | 86831 | 214539 | 418 | <LOD | 24004 | 1894 | 5361 | 193 | 55 | 1696 | 86980 | 299 | 50 | 49 | 135 | 25 | 6 | 189 | 49 | 37 | 151 | 15 | <LOD | 43 | 49 | 53 | 116 | 1348 | <LOD | <LOD | <LOD | 10 | <LOD | 10 | <LOD | 718231
10 [Col] | 2024-05-14 | NIST 679 | std | rep 5 | 9769 | 86280 | 213695 | 427 | <LOD | 23776 | 1941 | 4949 | 244 | 92 | 1579 | 85726 | 302 | 48 | 69 | 114 | 14 | 4 | 176 | 42 | 46 | 159 | 19 | 11 | <LOD | 24 | 83 | 46 | 886 | <LOD | <LOD | <LOD | 12 | <LOD | 12 | <LOD | 721842
11 [Col] | 2024-05-14 | SARM 69 | std | rep 1 | 9769 | 71191 | 250042 | 854 | <LOD | 17196 | 15037 | 4521 | 108 | 205 | 857 | 49523 | <LOD | 60 | 62 | 69 | 10 | 4 | 77 | 95 | 36 | 320 | 11 | 6 | 16 | 32 | 42 | 49 | 597 | <LOD | <LOD | <LOD | 14 | <LOD | 7 | <LOD | 733530
12 [Col] | 2024-05-14 | SARM 69 | std | rep 2 | 9769 | 70552 | 248683 | 821 | <LOD | 17164 | 14365 | 4602 | <LOD | 195 | 929 | 48913 | <LOD | 57 | 76 | 65 | 5 | 3 | 73 | 88 | 30 | 279 | 7 | 9 | <LOD | <LOD | 58 | 69 | 1064 | 20 | <LOD | <LOD | 13 | 23 | <LOD | 4 | 735887
13 [Col] | 2024-05-14 | SARM 69 | std | rep 3 | 13439 | 71516 | 257006 | 882 | <LOD | 17682 | 15379 | 4552 | 168 | 152 | 885 | 50340 | 126 | 56 | 68 | 72 | 8 | 3 | 78 | 111 | 31 | 330 | 8 | 10 | 31 | 31 | 38 | 56 | 686 | 25 | <LOD | <LOD | 18 | <LOD | 9 | <LOD | 720882
14 [Col] | 2024-05-14 | SARM 69 | std | rep 4 | 9769 | 69959 | 252919 | 865 | <LOD | 17135 | 15238 | 4527 | 202 | 161 | 904 | 49328 | 228 | 35 | 91 | 71 | 7 | 3 | 78 | 97 | 32 | 303 | 10 | 14 | <LOD | 36 | 70 | 87 | <LOD | 25 | <LOD | <LOD | 15 | <LOD | <LOD | <LOD | 732135
15 [Col] | 2024-05-14 | SARM 69 | std | rep 5 | 9769 | 69631 | 253411 | 786 | <LOD | 17271 | 14651 | 4590 | 240 | 156 | 873 | 49049 | 184 | 53 | 73 | 75 | 5 | 6 | 71 | 91 | 30 | 264 | 8 | 8 | 14 | 34 | 57 | 69 | 621 | 33 | <LOD | <LOD | 12 | <LOD | <LOD | <LOD | 732008

## averages
Test Label | Date | sample | component | rep | Mg | Al | Si | P | S | K | Ca | Ti | V | Cr | Mn | Fe | Co | Ni | Cu | Zn | As | Se | Rb | Sr | Y | Zr | Nb | Mo | Ag | Cd | Sn | Sb | Ba | W | Au | Hg | Pb | Bi | Th | U | LE
LLOD |  |  |  | Mg: 9769; Al: 22792; Si: 129981; K: 3448; Ca: 1140; Ti: 3031; Mn: 28; Fe: 991; Ni: 15; Zn: 5; Rb: 21; Sr: 23; Y: 5; Zr: 16; Nb: 5; Pb: 15
1 [Col] | 2023-06-06 | NIST 2711a | std | rep 1 | 15406 | 68395 | 261002 | 181 | 168 | 25265 | 23038 | <LOD | 78 | 48 | 656 | 29434 | <LOD | 27 | 710 | 378 | 150 | 7 | 119 | 268 | 31 | 281 | 18 | 23 | 40 | 105 | 49 | 125 | 1076 | <LOD | <LOD | 10 | 1378 | 24 | <LOD | <LOD | 728608
2 [Col] | 2023-06-06 | NIST 2711a | std | rep 2 | 15196 | 68055 | 257114 | 177 | 246 | 25100 | 22126 | 3961 | <LOD | 39 | 742 | 29897 | <LOD | 33 | 712 | 387 | 135 | <LOD | 123 | 263 | 29 | 263 | 16 | 7 | 16 | 76 | 42 | 57 | 1067 | <LOD | 6 | 20 | 1359 | <LOD | 18 | <LOD | 731904
3 [Col] | 2023-06-06 | NIST 2711a | std | rep 3 | 16362 | 68292 | 261493 | 85 | 245 | 24998 | 22349 | 3169 | 142 | 63 | 657 | 29356 | 148 | 42 | 699 | 383 | 74 | 4 | 124 | 262 | 29 | 306 | 24 | <LOD | 11 | 78 | 82 | 107 | 1201 | <LOD | 5 | 19 | 1407 | 33 | <LOD | <LOD | 727237
4 [Col] | 2023-06-06 | NIST 2711a | std | rep 4 | 14584 | 68798 | 258313 | 243 | 221 | 25170 | 22990 | 3126 | 186 | 95 | 684 | 28880 | <LOD | 36 | 727 | 341 | 102 | 2 | 121 | 259 | 39 | 310 | 24 | 13 | 14 | 74 | 53 | 132 | 655 | 44 | <LOD | 22 | 1347 | <LOD | 22 | <LOD | 732344
5 [Col] | 2023-06-06 | NIST 2711a | std | rep 5 | 16346 | 68521 | 262451 | 176 | 257 | 25769 | 22725 | <LOD | <LOD | 97 | 651 | 29052 | 122 | 21 | 718 | 398 | 145 | 3 | 120 | 257 | 32 | 294 | 13 | 4 | 39 | 89 | 76 | 140 | 1241 | 17 | <LOD | 14 | 1353 | <LOD | 11 | 4 | 726108
|  | NIST 2711a | 2023-06-06 | average | 15578.8 | 68412.2 | 260074.6 | 172.4 | 227.4 | 25260.4 | 22645.6 | 3418.6667 | 135.3333 | 68.4 | 678 | 29323.8 | <LOD | 31.8 | 713.2 | 377.4 | 121.2 | 4 | 121.4 | 261.8 | 32 | 290.8 | 19 | 11.75 | 24 | 84.4 | 60.4 | 112.2 | 1048 | <LOD | <LOD | 17 | 1368.8 | <LOD | <LOD | <LOD | 729240.2
|  | NIST 2711a | 2023-06-06 | stdev | 769.4187 | 275.2339 | 2257.6311 | 56.4163 | 35.7113 | 300.6032 | 399.0142 | 470.1663 | 54.3078 | 26.6233 | 38.0329 | 391.3352 | <LOD | 8.1056 | 10.3296 | 21.6402 | 32.3373 | 2.1602 | 2.0736 | 4.2071 | 4.1231 | 19.2276 | 4.899 | 8.3815 | 14.2653 | 12.8957 | 17.5585 | 33.1768 | 232.5145 | <LOD | <LOD | 4.899 | 24.3146 | <LOD | <LOD | <LOD | 2781.7516
|  | NIST 2711a | 2023-06-06 | RSD | 4.939% | 0.4023% | 0.8681% | 32.72% | 15.7% | 1.19% | 1.762% | 13.75% | 40.13% | 38.92% | 5.61% | 1.335% | <LOD | 25.49% | 1.448% | 5.734% | 26.68% | 54.01% | 1.708% | 1.607% | 12.88% | 6.612% | 25.78% | 71.33% | 59.44% | 15.28% | 29.07% | 29.57% | 22.19% | <LOD | <LOD | 28.82% | 1.776% | <LOD | <LOD | <LOD | 0.3815%
6 [Col] | 2023-06-06 | NIST 679 | std | rep 1 | 11618 | 90098 | 218211 | <LOD | <LOD | 23642 | 1554 | 5272 | 213 | 104 | 1353 | 91626 | 218 | 61 | 908 | 108 | 24 | 10 | 178 | 51 | 47 | 170 | 20 | 10 | 43 | 76 | 101 | 126 | 534 | 32 | <LOD | <LOD | <LOD | <LOD | <LOD | 16 | 706587
7 [Col] | 2023-06-06 | NIST 679 | std | rep 2 | 10844 | 90719 | 217962 | <LOD | <LOD | 23408 | 1303 | 5395 | 240 | 129 | 1561 | 89718 | 470 | 41 | 911 | 101 | 19 | 6 | 186 | 45 | 45 | 157 | 20 | 13 | 19 | 50 | <LOD | 49 | <LOD | <LOD | 9 | 8 | 20 | <LOD | 30 | <LOD | 709596
8 [Col] | 2023-06-06 | NIST 679 | std | rep 3 | 11975 | 90447 | 217339 | <LOD | <LOD | 23538 | 1495 | 5457 | 158 | 78 | 1457 | 88786 | 385 | 56 | 952 | 132 | 17 | 6 | 192 | 54 | 36 | 154 | 19 | 14 | 19 | 27 | 26 | 71 | 504 | 17 | <LOD | 6 | 17 | <LOD | 20 | <LOD | 709443
9 [Col] | 2023-06-06 | NIST 679 | std | rep 4 | 11771 | 87623 | 213484 | <LOD | <LOD | 23108 | 1492 | 4604 | 139 | 31 | 1521 | 90073 | 195 | 81 | 999 | 108 | 28 | 6 | 196 | 44 | 45 | 166 | 20 | 9 | 31 | <LOD | 22 | <LOD | 606 | <LOD | <LOD | 7 | <LOD | <LOD | 25 | 11 | 715442
10 [Col] | 2023-06-06 | NIST 679 | std | rep 5 | 12112 | 90489 | 217735 | <LOD | <LOD | 23431 | 1349 | 5363 | 138 | 107 | 1393 | 90512 | 268 | 55 | 900 | 109 | 21 | 2 | 194 | 53 | 43 | 154 | 17 | 12 | 35 | 30 | 59 | 90 | 1030 | 57 | <LOD | 8 | 16 | <LOD | 34 | <LOD | 707007
|  | NIST 679 | 2023-06-06 | average | 11664 | 89875.2 | 216946.2 | #DIV/0! | #DIV/0! | 23425.4 | 1438.6 | 5218.2 | 177.6 | 89.8 | 1457 | 90143 | 307.2 | 58.8 | 934 | 111.6 | 21.8 | 6 | 189.2 | 49.4 | 43.2 | 160.2 | 19.2 | 11.6 | 29.4 | 45.75 | 52 | 84 | 668.5 | 35.3333 | <LOD | 7.25 | 17.6667 | <LOD | 27.25 | <LOD | 709615
|  | NIST 679 | 2023-06-06 | stdev | 495.8301 | 1278.4511 | 1961.8164 | #DIV/0! | #DIV/0! | 200.4191 | 106.964 | 349.7809 | 46.339 | 37.5193 | 86.3481 | 1044.3208 | 116.8747 | 14.4637 | 41.5632 | 11.8448 | 4.3243 | 2.8284 | 7.2938 | 4.6152 | 4.2661 | 7.3621 | 1.3038 | 2.0736 | 10.4307 | 22.6035 | 36.6333 | 32.6292 | 244.7713 | 20.2073 | <LOD | 0.95743 | 2.0817 | <LOD | 6.0759 | <LOD | 3533.9186
|  | NIST 679 | 2023-06-06 | RSD | 4.251% | 1.422% | 0.9043% | #DIV/0! | #DIV/0! | 0.8556% | 7.435% | 6.703% | 26.09% | 41.78% | 5.926% | 1.159% | 38.05% | 24.6% | 4.45% | 10.61% | 19.84% | 47.14% | 3.855% | 9.342% | 9.875% | 4.596% | 6.791% | 17.88% | 35.48% | 49.41% | 70.45% | 38.84% | 36.62% | 57.19% | <LOD | 13.21% | 11.78% | <LOD | 22.3% | <LOD | 0.498%
11 [Col] | 2023-06-06 | SARM 69 | std | rep 1 | 13709 | 76526 | 257506 | 327 | <LOD | 17075 | 14308 | 4872 | 89 | 202 | 922 | 50482 | 183 | 77 | 720 | 62 | 12 | 3 | 72 | 110 | 29 | 286 | 7 | <LOD | 18 | 15 | 35 | 72 | 1258 | 15 | 5 | 9 | <LOD | <LOD | 11 | <LOD | 716059
12 [Col] | 2023-06-06 | SARM 69 | std | rep 2 | 13622 | 75672 | 256237 | 275 | <LOD | 17090 | 14055 | 4582 | 157 | 230 | 950 | 50338 | 184 | 82 | 717 | 70 | 6 | 2 | 73 | 106 | 30 | 258 | <LOD | <LOD | 32 | 26 | 42 | 74 | 808 | <LOD | 5 | <LOD | <LOD | 23 | <LOD | 6 | 718977
13 [Col] | 2023-06-06 | SARM 69 | std | rep 3 | 13780 | 78657 | 263033 | 378 | <LOD | 17728 | 14219 | 4894 | 73 | 230 | 892 | 51263 | 81 | 60 | 695 | 72 | 5 | 2 | 76 | 108 | 35 | 295 | 8 | 8 | 14 | 18 | 39 | 102 | 1016 | <LOD | 8 | 13 | 22 | <LOD | 30 | <LOD | 708343
14 [Col] | 2023-06-06 | SARM 69 | std | rep 4 | 14013 | 77359 | 259477 | 380 | <LOD | 17389 | 13905 | 4856 | 278 | 235 | 900 | 51579 | 66 | 67 | 680 | 58 | 6 | 3 | 76 | 100 | 29 | 338 | 16 | 11 | 18 | <LOD | 29 | 68 | 1351 | 16 | <LOD | 5 | 17 | <LOD | 11 | <LOD | 712015
15 [Col] | 2023-06-06 | SARM 69 | std | rep 5 | 14286 | 77760 | 260453 | 313 | <LOD | 17061 | 13797 | 4280 | 145 | 197 | 979 | 50676 | 258 | 92 | 656 | 70 | 6 | 3 | 69 | 97 | 30 | 271 | 14 | 16 | 37 | 35 | 31 | 72 | 2076 | <LOD | 6 | 6 | 18 | <LOD | 12 | 7 | 711496
|  | SARM 69 | 2023-06-06 | average | 13882 | 77194.8 | 259341.2 | 334.6 | <LOD | 17268.6 | 14056.8 | 4696.8 | 148.4 | 218.8 | 928.6 | 50867.6 | 154.4 | 75.6 | 693.6 | 66.4 | 7 | 2.6 | 73.2 | 104.2 | 30.6 | 289.6 | 11.25 | 11.6667 | 23.8 | 23.5 | 35.2 | 77.6 | 1301.8 | <LOD | 6 | 8.25 | 19 | <LOD | 16 | <LOD | 713378
|  | SARM 69 | 2023-06-06 | stdev | 268.4725 | 1145.7154 | 2640.4519 | 44.7806 | <LOD | 290.698 | 212.1042 | 265.4378 | 80.77 | 17.8241 | 36.0527 | 531.1086 | 80.0456 | 12.5419 | 26.6702 | 6.0663 | 2.8284 | 0.54772 | 2.9496 | 5.4955 | 2.51 | 30.5336 | 4.4253 | 4.0415 | 10.0598 | 8.9629 | 5.4037 | 13.813 | 481.9867 | <LOD | 1.4142 | 3.594 | 2.6458 | <LOD | 9.3452 | <LOD | 4161.9551
|  | SARM 69 | 2023-06-06 | RSD | 1.934% | 1.484% | 1.018% | 13.38% | <LOD | 1.683% | 1.509% | 5.651% | 54.43% | 8.146% | 3.882% | 1.044% | 51.84% | 16.59% | 3.845% | 9.136% | 40.41% | 21.07% | 4.029% | 5.274% | 8.203% | 10.54% | 39.34% | 34.64% | 42.27% | 38.14% | 15.35% | 17.8% | 37.02% | <LOD | 23.57% | 43.56% | 13.93% | <LOD | 58.41% | <LOD | 0.5834%
1 [Col] | 2023-06-07 | NIST 2711a | std | rep 1 | 14252 | 68263 | 262212 | 119 | 130 | 25592 | 23173 | 3722 | <LOD | 59 | 781 | 29821 | 188 | 38 | 698 | 399 | 142 | 3 | 135 | 266 | 31 | 312 | 25 | <LOD | 28 | 81 | 30 | 122 | 889 | <LOD | 5 | 22 | 1438 | <LOD | 14 | <LOD | 727393
2 [Col] | 2023-06-07 | NIST 2711a | std | rep 2 | 15171 | 68069 | 258473 | 196 | 152 | 24445 | 21616 | 3077 | 144 | 101 | 708 | 28677 | <LOD | 33 | 692 | 360 | 146 | 4 | 118 | 243 | 32 | 305 | 24 | 9 | <LOD | 68 | 44 | 44 | 1507 | 21 | 6 | <LOD | 1259 | 18 | <LOD | 10 | 733148
3 [Col] | 2023-06-07 | NIST 2711a | std | rep 3 | 15056 | 69390 | 258678 | 223 | 170 | 25575 | 22951 | 3202 | 103 | 88 | 675 | 28611 | 63 | 31 | 734 | 378 | 115 | <LOD | 119 | 248 | 34 | 324 | 16 | 9 | 19 | 77 | 57 | 68 | 1134 | <LOD | 9 | 14 | 1350 | <LOD | 23 | <LOD | 730645
4 [Col] | 2023-06-07 | NIST 2711a | std | rep 4 | 14898 | 69475 | 262233 | 81 | 123 | 25742 | 22588 | 3045 | 117 | 122 | 687 | 29967 | 66 | 28 | 716 | 411 | 133 | 4 | 122 | 271 | 24 | 296 | 24 | 7 | 23 | 104 | 33 | 74 | 662 | <LOD | 9 | 10 | 1382 | <LOD | 12 | <LOD | 726953
5 [Col] | 2023-06-07 | NIST 2711a | std | rep 5 | 15312 | 68777 | 258201 | 202 | 202 | 25197 | 22394 | 3651 | 150 | 80 | 765 | 28090 | 140 | 36 | 707 | 402 | 128 | 3 | 119 | 260 | 34 | 349 | 16 | 17 | 20 | 56 | 40 | 84 | 1111 | <LOD | 11 | 9 | 1313 | <LOD | 13 | <LOD | 731690
|  | NIST 2711a | 2023-06-07 | average | 14937.8 | 68794.8 | 259959.4 | 164.2 | 155.4 | 25310.2 | 22544.4 | 3339.4 | 128.5 | 90 | 723.2 | 29033.2 | 114.25 | 33.2 | 709.4 | 390 | 132.8 | 3.5 | 122.6 | 257.6 | 31 | 317.2 | 21 | 10.5 | 22.5 | 77.2 | 40.8 | 78.4 | 1060.6 | <LOD | 8 | 13.75 | 1348.4 | <LOD | 15.5 | <LOD | 729965.8
|  | NIST 2711a | 2023-06-07 | stdev | 412.3969 | 637.7399 | 2072.8486 | 60.9565 | 31.9812 | 523.7764 | 601.4377 | 323.2187 | 22.2486 | 23.5053 | 47.3096 | 819.6598 | 60.709 | 3.9623 | 16.4864 | 20.6761 | 12.2352 | 0.57735 | 7.0922 | 11.8448 | 4.1231 | 20.511 | 4.5826 | 4.4347 | 4.0415 | 17.7961 | 10.616 | 28.4746 | 314.4556 | <LOD | 2.4495 | 5.909 | 67.8329 | <LOD | 5.0662 | <LOD | 2704.4796
|  | NIST 2711a | 2023-06-07 | RSD | 2.761% | 0.927% | 0.7974% | 37.12% | 20.58% | 2.069% | 2.668% | 9.679% | 17.31% | 26.12% | 6.542% | 2.823% | 53.14% | 11.93% | 2.324% | 5.302% | 9.213% | 16.5% | 5.785% | 4.598% | 13.3% | 6.466% | 21.82% | 42.24% | 17.96% | 23.05% | 26.02% | 36.32% | 29.65% | <LOD | 30.62% | 42.97% | 5.031% | <LOD | 32.69% | <LOD | 0.3705%
6 [Col] | 2023-06-07 | NIST 679 | std | rep 1 | 10864 | 93377 | 222631 | <LOD | <LOD | 24106 | 1434 | 5696 | 218 | 121 | 1535 | 91622 | 236 | 45 | 844 | 117 | 20 | 10 | 187 | 51 | 42 | 163 | 18 | 10 | 32 | 43 | 58 | 52 | 698 | 18 | <LOD | <LOD | <LOD | <LOD | 16 | 9 | 700035
7 [Col] | 2023-06-07 | NIST 679 | std | rep 2 | 12601 | 90330 | 217378 | <LOD | <LOD | 23658 | 1471 | 5320 | 118 | 87 | 1640 | 89016 | 241 | 53 | 925 | 135 | 31 | 8 | 185 | 49 | 33 | 159 | 28 | 26 | <LOD | 16 | 57 | 64 | 630 | 21 | 4 | 10 | <LOD | <LOD | 15 | <LOD | 708503
8 [Col] | 2023-06-07 | NIST 679 | std | rep 3 | 10807 | 90790 | 218912 | <LOD | <LOD | 23302 | 1411 | 4882 | 177 | 92 | 1603 | 93207 | <LOD | 68 | 843 | 124 | 16 | 6 | 194 | 61 | 37 | 169 | 24 | 16 | <LOD | 21 | 33 | 73 | 1200 | <LOD | 5 | 12 | <LOD | <LOD | 29 | <LOD | 705051
9 [Col] | 2023-06-07 | NIST 679 | std | rep 4 | 11346 | 87036 | 210294 | <LOD | <LOD | 22146 | 1382 | 5142 | 116 | 167 | 1368 | 82667 | 297 | 40 | 762 | 117 | 23 | 4 | 167 | 44 | 43 | 151 | 14 | <LOD | 50 | 68 | 42 | 40 | 617 | <LOD | 6 | <LOD | <LOD | 41 | <LOD | <LOD | 726956
10 [Col] | 2023-06-07 | NIST 679 | std | rep 5 | 11895 | 90440 | 219276 | <LOD | <LOD | 23551 | 1400 | 4952 | 146 | 145 | 1733 | 93563 | 148 | 66 | 894 | 135 | 26 | 2 | 197 | 61 | 39 | 164 | 22 | 12 | <LOD | <LOD | 34 | 79 | 1423 | 30 | <LOD | 8 | <LOD | <LOD | 12 | <LOD | 702574
|  | NIST 679 | 2023-06-07 | average | 11502.6 | 90394.6 | 217698.2 | <LOD | <LOD | 23352.6 | 1419.6 | 5198.4 | 155 | 122.4 | 1575.8 | 90015 | 230.5 | 54.4 | 853.6 | 125.6 | 23.2 | 6 | 186 | 53.2 | 38.8 | 161.2 | 21.2 | 16 | <LOD | 37 | 44.8 | 61.6 | 913.6 | 23 | 5 | 10 | <LOD | <LOD | 18 | <LOD | 708623.8
|  | NIST 679 | 2023-06-07 | stdev | 754.4291 | 2254.8547 | 4561.2494 | <LOD | <LOD | 734.6236 | 34.3555 | 326.4947 | 43.0813 | 34.2023 | 136.399 | 4481.9282 | 61.5603 | 12.4218 | 61.8975 | 9.0443 | 5.7184 | 3.1623 | 11.7047 | 7.5631 | 4.0249 | 6.7231 | 5.4037 | 7.1181 | <LOD | 23.7627 | 12.112 | 15.7892 | 372.9602 | 6.245 | 1 | 2 | <LOD | <LOD | 7.5277 | <LOD | 10714.67
|  | NIST 679 | 2023-06-07 | RSD | 6.559% | 2.494% | 2.095% | <LOD | <LOD | 3.146% | 2.42% | 6.281% | 27.79% | 27.94% | 8.656% | 4.979% | 26.71% | 22.83% | 7.251% | 7.201% | 24.65% | 52.7% | 6.293% | 14.22% | 10.37% | 4.171% | 25.49% | 44.49% | <LOD | 64.22% | 27.04% | 25.63% | 40.82% | 27.15% | 20% | 20% | <LOD | <LOD | 41.82% | <LOD | 1.512%
11 [Col] | 2023-06-07 | SARM 69 | std | rep 1 | 13621 | 77075 | 261201 | 392 | <LOD | 17468 | 13843 | 4173 | 298 | 203 | 909 | 50482 | 70 | 48 | 653 | 51 | 11 | 3 | 73 | 96 | 28 | 347 | 16 | 6 | 16 | 33 | 58 | 60 | 470 | 72 | <LOD | 7 | <LOD | 44 | <LOD | 8 | 713708
12 [Col] | 2023-06-07 | SARM 69 | std | rep 2 | 13424 | 76567 | 262417 | 275 | <LOD | 17404 | 13992 | 4667 | 95 | 227 | 978 | 51539 | 230 | 82 | 639 | 63 | 10 | 4 | 79 | 102 | 25 | 247 | 9 | <LOD | <LOD | 35 | 63 | 46 | 799 | 29 | <LOD | 6 | <LOD | <LOD | 8 | <LOD | 711467
13 [Col] | 2023-06-07 | SARM 69 | std | rep 3 | 13145 | 78056 | 261324 | 323 | <LOD | 17660 | 14190 | 4728 | 132 | 250 | 1004 | 51741 | 182 | 49 | 667 | 59 | 8 | 7 | 79 | 94 | 30 | 272 | 9 | 8 | 27 | 14 | 60 | 96 | 1262 | <LOD | 4 | <LOD | 17 | <LOD | <LOD | <LOD | 710513
14 [Col] | 2023-06-07 | SARM 69 | std | rep 4 | 13194 | 78612 | 266546 | 328 | <LOD | 17039 | 14030 | 4593 | 136 | 275 | 952 | 50942 | 312 | 74 | 615 | 71 | 11 | 5 | 73 | 103 | 28 | 282 | 10 | <LOD | 23 | 47 | 70 | 34 | 605 | 30 | <LOD | <LOD | 16 | <LOD | 16 | <LOD | 707152
15 [Col] | 2023-06-07 | SARM 69 | std | rep 5 | 11923 | 78182 | 265931 | 366 | <LOD | 17292 | 14315 | 4276 | 173 | 262 | 985 | 51250 | 162 | 46 | 663 | 71 | 9 | 7 | 75 | 110 | 27 | 293 | 8 | 7 | 36 | 58 | 55 | 51 | 1531 | <LOD | <LOD | 6 | <LOD | <LOD | 14 | 7 | 708323
|  | SARM 69 | 2023-06-07 | average | 13061.4 | 77698.4 | 263483.8 | 336.8 | <LOD | 17372.6 | 14074 | 4487.4 | 166.8 | 243.4 | 965.6 | 51190.8 | 191.2 | 59.8 | 647.4 | 63 | 9.8 | 5.2 | 75.8 | 101 | 27.6 | 288.2 | 10.4 | 7 | 25.5 | 37.4 | 61.2 | 57.4 | 933.4 | 43.6667 | <LOD | 6.3333 | <LOD | <LOD | 12.6667 | <LOD | 710232.6
|  | SARM 69 | 2023-06-07 | stdev | 664.2961 | 846.3311 | 2568.065 | 44.6844 | <LOD | 229.3857 | 182.7279 | 247.4031 | 78.369 | 28.6758 | 36.7192 | 497.9374 | 89.0573 | 16.8879 | 21.0903 | 8.4853 | 1.3038 | 1.7889 | 3.0332 | 6.3246 | 1.8166 | 37.0095 | 3.2094 | 1 | 8.3467 | 16.5015 | 5.7184 | 23.533 | 448.8188 | 24.5425 | <LOD | 0.57735 | <LOD | <LOD | 4.1633 | <LOD | 2589.2779
|  | SARM 69 | 2023-06-07 | RSD | 5.086% | 1.089% | 0.9747% | 13.27% | <LOD | 1.32% | 1.298% | 5.513% | 46.98% | 11.78% | 3.803% | 0.9727% | 46.58% | 28.24% | 3.258% | 13.47% | 13.3% | 34.4% | 4.002% | 6.262% | 6.582% | 12.84% | 30.86% | 14.29% | 32.73% | 44.12% | 9.344% | 41% | 48.08% | 56.2% | <LOD | 9.116% | <LOD | <LOD | 32.87% | <LOD | 0.3646%
2 [Col] | 2023-06-08 | NIST 2711a | std | rep 1 | 14108 | 68765 | 261965 | 147 | 216 | 25611 | 22608 | 3145 | 85 | 38 | 655 | 28830 | 103 | 47 | 555 | 370 | 117 | 4 | 125 | 271 | 33 | 319 | 20 | 9 | <LOD | 76 | 74 | 98 | 1772 | 53 | <LOD | 6 | 1389 | <LOD | 9 | 5 | 728724
3 [Col] | 2023-06-08 | NIST 2711a | std | rep 2 | 14786 | 67043 | 258900 | 193 | 199 | 25394 | 22975 | 3231 | 135 | 42 | 749 | 28990 | 204 | 26 | 587 | 414 | 144 | <LOD | 124 | 249 | 31 | 314 | 20 | 5 | 23 | 59 | 37 | 37 | 1229 | 23 | <LOD | <LOD | 1416 | <LOD | 29 | <LOD | 732041
4 [Col] | 2023-06-08 | NIST 2711a | std | rep 3 | 15057 | 67779 | 261198 | 200 | 149 | 24613 | 22718 | 3586 | 119 | 143 | 654 | 29332 | 82 | 24 | 543 | 406 | 123 | 3 | 115 | 259 | 32 | 291 | 21 | 16 | 35 | 88 | 43 | 63 | 1029 | 28 | <LOD | 8 | 1310 | <LOD | 25 | <LOD | 729430
5 [Col] | 2023-06-08 | NIST 2711a | std | rep 4 | 15533 | 67327 | 255281 | 205 | 184 | 24654 | 22468 | 3363 | 75 | 126 | 561 | 28655 | <LOD | 30 | 568 | 380 | 125 | 3 | 115 | 258 | 32 | 287 | 21 | 9 | 22 | 95 | 70 | 85 | 1184 | 48 | <LOD | 8 | 1296 | 26 | <LOD | <LOD | 735741
6 [Col] | 2023-06-08 | NIST 2711a | std | rep 5 | 15071 | 68115 | 257607 | 184 | 161 | 24845 | 22090 | 3083 | 202 | 153 | 643 | 28079 | 137 | 43 | 546 | 404 | 149 | 3 | 116 | 245 | 33 | 315 | 17 | 10 | <LOD | 71 | 48 | 91 | 629 | 17 | <LOD | 8 | 1315 | <LOD | 17 | <LOD | 734761
|  | NIST 2711a | 2023-06-08 | average | 14911 | 67805.8 | 258990.2 | 185.8 | 181.8 | 25023.4 | 22571.8 | 3281.6 | 123.2 | 100.4 | 652.4 | 28777.2 | 131.5 | 34 | 559.8 | 394.8 | 131.6 | 3.25 | 119 | 256.4 | 32.2 | 305.2 | 19.8 | 9.8 | 26.6667 | 77.8 | 54.4 | 74.8 | 1168.6 | 33.8 | <LOD | 7.5 | 1345.2 | <LOD | 20 | <LOD | 732139.4
|  | NIST 2711a | 2023-06-08 | stdev | 523.0282 | 675.9254 | 2711.2432 | 23.0803 | 27.2892 | 452.584 | 327.17 | 199.967 | 50.3706 | 55.9937 | 66.6768 | 463.2199 | 53.3823 | 10.3682 | 18.0472 | 18.7936 | 14.0285 | 0.5 | 5.0498 | 10.0896 | 0.83666 | 14.9733 | 1.6432 | 3.9623 | 7.2342 | 14.1669 | 16.5922 | 24.8636 | 411.896 | 15.8335 | <LOD | 1 | 53.6256 | <LOD | 8.8694 | <LOD | 3116.8813
|  | NIST 2711a | 2023-06-08 | RSD | 3.508% | 0.9969% | 1.047% | 12.42% | 15.01% | 1.809% | 1.449% | 6.094% | 40.89% | 55.77% | 10.22% | 1.61% | 40.59% | 30.49% | 3.224% | 4.76% | 10.66% | 15.38% | 4.243% | 3.935% | 2.598% | 4.906% | 8.299% | 40.43% | 27.13% | 18.21% | 30.5% | 33.24% | 35.25% | 46.84% | <LOD | 13.33% | 3.986% | <LOD | 44.35% | <LOD | 0.4257%
7 [Col] | 2023-06-08 | NIST 679 | std | rep 1 | <LOD | 92172 | 219291 | <LOD | <LOD | 23535 | 1384 | 5409 | 220 | <LOD | 1603 | 89537 | 165 | 52 | 593 | 100 | 22 | 2 | 179 | 41 | 31 | 156 | 16 | 12 | 32 | 58 | 60 | 88 | 634 | 58 | <LOD | <LOD | <LOD | <LOD | 18 | <LOD | 708506
8 [Col] | 2023-06-08 | NIST 679 | std | rep 2 | 11661 | 94095 | 222594 | <LOD | <LOD | 24338 | 1427 | 5180 | 210 | 94 | 1676 | 91289 | 265 | 64 | 624 | 126 | 24 | 5 | 194 | 49 | 33 | 159 | 23 | 11 | 28 | 48 | 62 | 94 | <LOD | 62 | <LOD | <LOD | 17 | <LOD | 41 | <LOD | 699973
9 [Col] | 2023-06-08 | NIST 679 | std | rep 3 | 12091 | 94632 | 224001 | <LOD | <LOD | 24182 | 1487 | 5177 | 143 | 109 | 1733 | 92700 | 240 | 63 | 640 | 121 | 24 | 6 | 192 | 48 | 32 | 156 | 24 | <LOD | <LOD | 15 | 53 | 71 | <LOD | <LOD | 6 | <LOD | <LOD | <LOD | 10 | 5 | 696609
10 [Col] | 2023-06-08 | NIST 679 | std | rep 4 | 11492 | 92810 | 222277 | <LOD | <LOD | 23724 | 1308 | 5678 | 127 | 102 | 1565 | 91513 | 326 | 71 | 629 | 130 | 18 | 5 | 182 | 51 | 44 | 176 | 19 | 15 | 47 | 41 | 60 | 79 | 854 | 40 | <LOD | <LOD | 18 | <LOD | 24 | <LOD | 700429
11 [Col] | 2023-06-08 | NIST 679 | std | rep 5 | 12289 | 91531 | 220863 | <LOD | <LOD | 23832 | 1357 | 5010 | 160 | 118 | 1557 | 91485 | 183 | 50 | 598 | 118 | 23 | 7 | 198 | 53 | 39 | 163 | 11 | 14 | 40 | 68 | 60 | 164 | 605 | <LOD | <LOD | 7 | <LOD | <LOD | 36 | <LOD | 702800
|  | NIST 679 | 2023-06-08 | average | 11883.25 | 93048 | 221805.2 | <LOD | <LOD | 23922.2 | 1392.6 | 5290.8 | 172 | 105.75 | 1626.8 | 91304.8 | 235.8 | 60 | 616.8 | 119 | 22.2 | 5 | 189 | 48.4 | 35.8 | 162 | 18.6 | 13 | 36.75 | 46 | 59 | 99.2 | 697.6667 | 53.3333 | <LOD | <LOD | <LOD | <LOD | 25.8 | <LOD | 701663.4
|  | NIST 679 | 2023-06-08 | stdev | 369.8039 | 1297.1675 | 1794.1071 | <LOD | <LOD | 330.8054 | 68.1784 | 258.8295 | 41.1035 | 10.2103 | 75.7575 | 1134.4347 | 64.8205 | 8.8034 | 20.3642 | 11.5758 | 2.49 | 1.8708 | 8.124 | 4.5607 | 5.5408 | 8.3367 | 5.3198 | 1.8257 | 8.4607 | 20.1122 | 3.4641 | 37.2653 | 136.1629 | 11.7189 | <LOD | <LOD | <LOD | <LOD | 12.7358 | <LOD | 4417.0462
|  | NIST 679 | 2023-06-08 | RSD | 3.112% | 1.394% | 0.8089% | <LOD | <LOD | 1.383% | 4.896% | 4.892% | 23.9% | 9.655% | 4.657% | 1.242% | 27.49% | 14.67% | 3.302% | 9.728% | 11.22% | 37.42% | 4.298% | 9.423% | 15.48% | 5.146% | 28.6% | 14.04% | 23.02% | 43.72% | 5.871% | 37.57% | 19.52% | 21.97% | <LOD | <LOD | <LOD | <LOD | 49.36% | <LOD | 0.6295%
12 [Col] | 2023-06-08 | SARM 69 | std | rep 1 | 14791 | 78606 | 265484 | 425 | <LOD | 17717 | 13876 | 4504 | 314 | 199 | 1070 | 51427 | 235 | 51 | 518 | 58 | 5 | 2 | 77 | 106 | 33 | 321 | 12 | 7 | 18 | 25 | 55 | 80 | 878 | 39 | <LOD | <LOD | <LOD | <LOD | 9 | 3 | 705151
13 [Col] | 2023-06-08 | SARM 69 | std | rep 2 | 14941 | 76553 | 259836 | 330 | <LOD | 17282 | 13434 | 4876 | 212 | 245 | 826 | 51466 | 179 | 59 | 536 | 54 | 17 | 2 | 77 | 108 | 24 | 241 | 8 | 6 | 29 | 50 | 46 | <LOD | 857 | 25 | <LOD | <LOD | <LOD | <LOD | 21 | 5 | 712479
14 [Col] | 2023-06-08 | SARM 69 | std | rep 3 | 13904 | 79413 | 267378 | 365 | <LOD | 17716 | 14706 | 4364 | 196 | 207 | 922 | 51355 | 242 | 60 | 534 | 65 | 6 | 7 | 77 | 95 | 32 | 314 | 12 | 9 | 14 | 35 | 66 | 64 | 1029 | <LOD | 4 | 7 | <LOD | <LOD | 15 | <LOD | 703670
15 [Col] | 2023-06-08 | SARM 69 | std | rep 4 | 12686 | 76793 | 264770 | 344 | <LOD | 17444 | 14059 | 4701 | 309 | 176 | 1181 | 52423 | <LOD | 53 | 536 | 71 | 7 | 2 | 82 | 113 | 28 | 294 | 7 | 15 | <LOD | 73 | 58 | 72 | 806 | <LOD | 8 | 7 | <LOD | 16 | <LOD | <LOD | 708787
16 [Col] | 2023-06-08 | SARM 69 | std | rep 5 | 13693 | 78361 | 263921 | 369 | <LOD | 17205 | 13842 | 4304 | 198 | 147 | 889 | 50413 | <LOD | 46 | 492 | 78 | 4 | 9 | 78 | 102 | 27 | 254 | 8 | 6 | 32 | 46 | 66 | 47 | 622 | <LOD | <LOD | 10 | 25 | <LOD | <LOD | <LOD | 710613
|  | SARM 69 | 2023-06-08 | average | 14003 | 77945.2 | 264277.8 | 366.6 | <LOD | 17472.8 | 13983.4 | 4549.8 | 245.8 | 194.8 | 977.6 | 51416.8 | 218.6667 | 53.8 | 523.2 | 65.2 | 7.8 | 4.4 | 78.2 | 104.8 | 28.8 | 284.8 | 9.4 | 8.6 | 23.25 | 45.8 | 58.2 | 65.75 | 838.4 | <LOD | 6 | 8 | <LOD | <LOD | 15 | <LOD | 708140
|  | SARM 69 | 2023-06-08 | stdev | 913.9445 | 1227.7794 | 2791.2951 | 36.2946 | <LOD | 238.6057 | 464.0138 | 237.8575 | 60.3175 | 36.4856 | 144.7698 | 711.7578 | 34.5302 | 5.8052 | 19.0053 | 9.6799 | 5.2631 | 3.3615 | 2.1679 | 6.7602 | 3.7014 | 35.7589 | 2.4083 | 3.7815 | 8.6168 | 18.0748 | 8.3785 | 14.1038 | 146.7321 | <LOD | 2.8284 | 1.7321 | <LOD | <LOD | 6 | <LOD | 3683.6219
|  | SARM 69 | 2023-06-08 | RSD | 6.527% | 1.575% | 1.056% | 9.9% | <LOD | 1.366% | 3.318% | 5.228% | 24.54% | 18.73% | 14.81% | 1.384% | 15.79% | 10.79% | 3.633% | 14.85% | 67.48% | 76.4% | 2.772% | 6.451% | 12.85% | 12.56% | 25.62% | 43.97% | 37.06% | 39.46% | 14.4% | 21.45% | 17.5% | <LOD | 47.14% | 21.65% | <LOD | <LOD | 40% | <LOD | 0.5202%
1 [Col] | 2023-06-09 | NIST 2711a | std | rep 1 | 13434 | 67975 | 256138 | 203 | 267 | 24449 | 22510 | 3056 | 92 | 74 | 645 | 27423 | <LOD | 24 | 595 | 384 | 145 | <LOD | 116 | 244 | 29 | 277 | 21 | 9 | 21 | 87 | 51 | 83 | 1697 | <LOD | 7 | 13 | 1296 | <LOD | 13 | 8 | 737954
2 [Col] | 2023-06-09 | NIST 2711a | std | rep 2 | 15238 | 68918 | 259395 | 140 | 114 | 25333 | 22224 | 3895 | 113 | 33 | 692 | 28961 | 65 | 30 | 555 | 387 | 111 | 6 | 123 | 277 | 43 | 290 | 14 | 6 | 19 | 82 | 50 | 48 | 975 | <LOD | 6 | 10 | 1353 | <LOD | 8 | <LOD | 730173
3 [Col] | 2023-06-09 | NIST 2711a | std | rep 3 | 14197 | 67748 | 257255 | 190 | 132 | 24678 | 22375 | <LOD | 75 | 74 | 589 | 28691 | 65 | 24 | 559 | 396 | 146 | <LOD | 114 | 254 | 30 | 317 | 15 | 6 | 17 | 82 | 61 | 115 | 1457 | 38 | <LOD | 6 | 1317 | <LOD | <LOD | <LOD | 735564
4 [Col] | 2023-06-09 | NIST 2711a | std | rep 4 | 15624 | 70506 | 261592 | 166 | 214 | 25141 | 23285 | 3491 | <LOD | 165 | 696 | 29608 | 58 | 37 | 645 | 406 | 140 | 5 | 128 | 271 | 31 | 339 | 20 | <LOD | 15 | 106 | 31 | 79 | 1518 | <LOD | 10 | 20 | 1352 | <LOD | <LOD | <LOD | 724786
5 [Col] | 2023-06-09 | NIST 2711a | std | rep 5 | 16523 | 71054 | 262067 | 217 | 201 | 25713 | 22431 | 3233 | 66 | 124 | 715 | 29267 | 92 | 27 | 560 | 394 | 124 | 2 | 118 | 254 | 31 | 308 | 20 | 4 | 40 | 124 | 86 | 77 | 1404 | <LOD | <LOD | 12 | 1384 | <LOD | <LOD | <LOD | 723754
|  | NIST 2711a | 2023-06-09 | average | 15003.2 | 69240.2 | 259289.4 | 183.2 | 185.6 | 25062.8 | 22565 | 3418.75 | 86.5 | 94 | 667.4 | 28790 | 70 | 28.4 | 582.8 | 393.4 | 133.2 | 4.3333 | 119.8 | 260 | 32.8 | 306.2 | 18 | 6.25 | 22.4 | 96.2 | 55.8 | 80.4 | 1410.2 | <LOD | 7.6667 | 12.2 | 1340.4 | <LOD | <LOD | <LOD | 730446.2
|  | NIST 2711a | 2023-06-09 | stdev | 1210.6468 | 1485.2041 | 2602.8007 | 30.5892 | 62.5883 | 506.6322 | 415.8612 | 364.2914 | 20.6962 | 51.142 | 50.836 | 837.3148 | 15.0333 | 5.4129 | 38.3301 | 8.5907 | 15.2217 | 2.0817 | 5.6745 | 13.5831 | 5.7619 | 24.0354 | 3.2404 | 2.0616 | 10.0896 | 18.4174 | 20.0674 | 23.8076 | 267.1623 | <LOD | 2.0817 | 5.1186 | 34.3264 | <LOD | <LOD | <LOD | 6313.8138
|  | NIST 2711a | 2023-06-09 | RSD | 8.069% | 2.145% | 1.004% | 16.7% | 33.72% | 2.021% | 1.843% | 10.66% | 23.93% | 54.41% | 7.617% | 2.908% | 21.48% | 19.06% | 6.577% | 2.184% | 11.43% | 48.04% | 4.737% | 5.224% | 17.57% | 7.85% | 18% | 32.98% | 45.04% | 19.14% | 35.96% | 29.61% | 18.94% | <LOD | 27.15% | 41.96% | 2.561% | <LOD | <LOD | <LOD | 0.8644%
6 [Col] | 2023-06-09 | NIST 679 | std | rep 1 | <LOD | 91309 | 219204 | <LOD | <LOD | 23521 | 1449 | 4883 | 158 | 90 | 1533 | 88714 | <LOD | 66 | 593 | 109 | 19 | 7 | 191 | 50 | 35 | 156 | 27 | 21 | 29 | 20 | 27 | 83 | 796 | <LOD | 9 | <LOD | 20 | 22 | <LOD | <LOD | 710706
7 [Col] | 2023-06-09 | NIST 679 | std | rep 2 | 10882 | 91716 | 220485 | <LOD | <LOD | 24023 | 1501 | 5414 | 181 | 137 | 1648 | 92110 | <LOD | 48 | 652 | 133 | 37 | 7 | 191 | 59 | 43 | 163 | 24 | 23 | 15 | 43 | 62 | 63 | <LOD | 42 | <LOD | <LOD | <LOD | <LOD | 25 | <LOD | 704058
8 [Col] | 2023-06-09 | NIST 679 | std | rep 3 | 12055 | 94045 | 220807 | <LOD | <LOD | 23708 | 1427 | 5335 | 199 | 82 | 1653 | 91603 | 166 | 53 | 597 | 137 | 32 | 7 | 176 | 47 | 39 | 173 | 27 | <LOD | <LOD | 42 | 32 | 27 | 1230 | <LOD | <LOD | 6 | <LOD | <LOD | 12 | <LOD | 700255
9 [Col] | 2023-06-09 | NIST 679 | std | rep 4 | 11354 | 92525 | 223167 | <LOD | <LOD | 24288 | 1402 | 5109 | 125 | 131 | 1624 | 90911 | 261 | 51 | 621 | 116 | 14 | 3 | 204 | 56 | 38 | 164 | 25 | 11 | 55 | 60 | 71 | 95 | 708 | 33 | 4 | <LOD | <LOD | <LOD | 23 | <LOD | 700937
10 [Col] | 2023-06-09 | NIST 679 | std | rep 5 | 12477 | 88158 | 212333 | <LOD | <LOD | 22390 | 1347 | 5578 | 188 | 89 | 1518 | 84854 | <LOD | 50 | 605 | 140 | 22 | 9 | 163 | 37 | 41 | 138 | 18 | 4 | <LOD | 31 | 87 | 106 | 671 | <LOD | <LOD | 6 | <LOD | <LOD | <LOD | <LOD | 720337
|  | NIST 679 | 2023-06-09 | average | 11692 | 91550.6 | 219199.2 | <LOD | <LOD | 23586 | 1425.2 | 5263.8 | 170.2 | 105.8 | 1595.2 | 89638.4 | <LOD | 53.6 | 613.6 | 127 | 24.8 | 6.6 | 185 | 49.8 | 39.2 | 158.8 | 24.2 | 14.75 | 33 | 39.2 | 55.8 | 74.8 | 851.25 | <LOD | <LOD | <LOD | <LOD | <LOD | 20 | <LOD | 707258.6
|  | NIST 679 | 2023-06-09 | stdev | 711.416 | 2166.0419 | 4096.5371 | <LOD | <LOD | 730.2907 | 56.9403 | 271.7181 | 29.3888 | 26.0135 | 64.7819 | 2972.0771 | <LOD | 7.1624 | 23.9958 | 13.6931 | 9.471 | 2.1909 | 15.7956 | 8.5849 | 3.0332 | 13.1034 | 3.7014 | 8.8835 | 20.2978 | 14.9231 | 25.6846 | 31.1159 | 257.8855 | <LOD | <LOD | <LOD | <LOD | <LOD | 7 | <LOD | 8399.1963
|  | NIST 679 | 2023-06-09 | RSD | 6.085% | 2.366% | 1.869% | <LOD | <LOD | 3.096% | 3.995% | 5.162% | 17.27% | 24.59% | 4.061% | 3.316% | <LOD | 13.36% | 3.911% | 10.78% | 38.19% | 33.2% | 8.538% | 17.24% | 7.738% | 8.252% | 15.29% | 60.23% | 61.51% | 38.07% | 46.03% | 41.6% | 30.29% | <LOD | <LOD | <LOD | <LOD | <LOD | 35% | <LOD | 1.188%
11 [Col] | 2023-06-09 | SARM 69 | std | rep 1 | 12303 | 76517 | 261536 | 293 | <LOD | 17335 | 13750 | 4693 | <LOD | 174 | 921 | 51086 | 130 | 53 | 523 | 73 | 7 | 4 | 84 | 114 | 37 | 290 | 11 | 18 | <LOD | 24 | 60 | 81 | 1198 | 19 | <LOD | <LOD | 18 | <LOD | 15 | <LOD | 714169
12 [Col] | 2023-06-09 | SARM 69 | std | rep 2 | 13268 | 77329 | 259925 | 366 | <LOD | 17340 | 13937 | 5219 | <LOD | 252 | 875 | 50994 | 242 | 78 | 528 | 49 | 5 | 4 | 72 | 103 | 33 | 322 | 12 | 7 | 18 | 33 | 66 | 63 | 658 | 21 | <LOD | 8 | 21 | 21 | <LOD | <LOD | 713600
13 [Col] | 2023-06-09 | SARM 69 | std | rep 3 | 12999 | 78872 | 264625 | 303 | <LOD | 17928 | 14099 | 4750 | <LOD | 190 | 917 | 52254 | 173 | 72 | 494 | 68 | 6 | 4 | 75 | 104 | 35 | 314 | 10 | <LOD | 12 | <LOD | 35 | 51 | 1003 | <LOD | 8 | <LOD | 22 | <LOD | 18 | 8 | 707092
14 [Col] | 2023-06-09 | SARM 69 | std | rep 4 | 12894 | 77526 | 263657 | 319 | <LOD | 17470 | 14275 | 4881 | <LOD | 248 | 853 | 51818 | 99 | 51 | 538 | 81 | 8 | 3 | 78 | 105 | 35 | 324 | 6 | 7 | <LOD | 29 | 71 | 58 | <LOD | <LOD | 4 | <LOD | <LOD | <LOD | 10 | <LOD | 710595
15 [Col] | 2023-06-09 | SARM 69 | std | rep 5 | 12982 | 78326 | 264064 | 295 | <LOD | 17553 | 13880 | 5213 | 123 | 198 | 850 | 51862 | 68 | 94 | 544 | 62 | 7 | <LOD | 77 | 105 | 31 | 334 | 9 | 23 | <LOD | <LOD | 36 | 53 | 877 | 30 | <LOD | 9 | 19 | <LOD | 12 | <LOD | 708375
|  | SARM 69 | 2023-06-09 | average | 12889.2 | 77714 | 262761.4 | 315.2 | <LOD | 17525.2 | 13988.2 | 4951.2 | <LOD | 212.4 | 883.2 | 51602.8 | 142.4 | 69.6 | 525.4 | 66.6 | 6.6 | 3.75 | 77.2 | 106.2 | 34.2 | 316.8 | 9.6 | 13.75 | <LOD | 28.6667 | 53.6 | 61.2 | 934 | 23.3333 | <LOD | <LOD | 20 | <LOD | 13.75 | <LOD | 710766.2
|  | SARM 69 | 2023-06-09 | stdev | 356.3042 | 912.6727 | 1969.6853 | 30.1861 | <LOD | 243.1783 | 203.4692 | 251.1637 | <LOD | 35.4232 | 34.1057 | 542.0509 | 67.8771 | 17.9805 | 19.3856 | 12.054 | 1.1402 | 0.5 | 4.4385 | 4.4385 | 2.2804 | 16.5892 | 2.3022 | 8.0571 | <LOD | 4.5092 | 16.9794 | 12.0083 | 226.4818 | 5.8595 | <LOD | <LOD | 1.8257 | <LOD | 3.5 | <LOD | 3116.7436
|  | SARM 69 | 2023-06-09 | RSD | 2.764% | 1.174% | 0.7496% | 9.577% | <LOD | 1.388% | 1.455% | 5.073% | <LOD | 16.68% | 3.862% | 1.05% | 47.67% | 25.83% | 3.69% | 18.1% | 17.28% | 13.33% | 5.749% | 4.179% | 6.668% | 5.236% | 23.98% | 58.6% | <LOD | 15.73% | 31.68% | 19.62% | 24.25% | 25.11% | <LOD | <LOD | 9.129% | <LOD | 25.45% | <LOD | 0.4385%
1 [Col] | 2023-06-13 | NIST 2711a | std | rep 1 | 15464 | 72482 | 266609 | 276 | 189 | 26199 | 23712 | 3797 | 204 | 48 | 657 | 29081 | 91 | 37 | 443 | 403 | 119 | 3 | 119 | 269 | 35 | 342 | 20 | 5 | <LOD | 72 | 76 | 86 | 769 | <LOD | <LOD | 6 | 1362 | <LOD | 7 | 3 | 718921
2 [Col] | 2023-06-13 | NIST 2711a | std | rep 2 | 15788 | 70680 | 266770 | 121 | 114 | 25453 | 22690 | 3156 | 91 | 133 | 593 | 28839 | 76 | 32 | 378 | 368 | 138 | <LOD | 122 | 256 | 33 | 297 | 18 | 5 | 40 | 101 | 77 | 99 | 913 | <LOD | <LOD | 19 | 1286 | <LOD | 23 | 5 | 722130
3 [Col] | 2023-06-13 | NIST 2711a | std | rep 3 | 16657 | 70577 | 264937 | 187 | 210 | 25445 | 23440 | 3754 | 131 | <LOD | 639 | 28768 | <LOD | 23 | 407 | 396 | 127 | 2 | 121 | 249 | 28 | 279 | 18 | 5 | 35 | 103 | 44 | 58 | 1202 | <LOD | 10 | 18 | 1329 | <LOD | <LOD | 6 | 721553
4 [Col] | 2023-06-13 | NIST 2711a | std | rep 4 | 15610 | 69751 | 265080 | 127 | 112 | 25590 | 22007 | <LOD | <LOD | 100 | 695 | 28978 | 71 | 20 | 335 | 386 | 154 | 3 | 124 | 258 | 28 | 310 | 17 | 14 | <LOD | 59 | 43 | 115 | 1436 | 16 | <LOD | 15 | 1323 | 33 | 8 | <LOD | 724535
5 [Col] | 2023-06-13 | NIST 2711a | std | rep 5 | 15249 | 70785 | 265107 | 178 | 229 | 25311 | 23296 | 3336 | <LOD | 105 | 688 | 28114 | 111 | 24 | 351 | 384 | 124 | 5 | 117 | 239 | 27 | 273 | 23 | 13 | <LOD | 49 | 62 | 90 | 1241 | <LOD | <LOD | 10 | 1324 | 19 | <LOD | 5 | 724083
|  | NIST 2711a | 2023-06-13 | average | 15753.6 | 70855 | 265700.6 | 177.8 | 170.8 | 25599.6 | 23029 | 3510.75 | 142 | 96.5 | 654.4 | 28756 | 87.25 | 27.2 | 382.8 | 387.4 | 132.4 | 3.25 | 120.6 | 254.2 | 30.2 | 300.2 | 19.2 | 8.4 | <LOD | 76.8 | 60.4 | 89.6 | 1112.2 | <LOD | <LOD | 13.6 | 1324.8 | <LOD | 12.6667 | 4.75 | 722244.4
|  | NIST 2711a | 2023-06-13 | stdev | 542.3157 | 997.3407 | 906.8348 | 62.3193 | 54.6324 | 349.3047 | 683.0271 | 314.9046 | 57.2975 | 35.4448 | 41.2044 | 378.8489 | 17.9699 | 7.0498 | 43.3843 | 13.2966 | 13.9392 | 1.2583 | 2.7019 | 11.1221 | 3.5637 | 27.5808 | 2.3875 | 4.669 | <LOD | 24.4172 | 16.5318 | 20.8878 | 267.8464 | <LOD | <LOD | 5.5045 | 26.9759 | <LOD | 8.9629 | 1.2583 | 2245.1841
|  | NIST 2711a | 2023-06-13 | RSD | 3.442% | 1.408% | 0.3413% | 35.05% | 31.99% | 1.364% | 2.966% | 8.97% | 40.35% | 36.73% | 6.297% | 1.317% | 20.6% | 25.92% | 11.33% | 3.432% | 10.53% | 38.72% | 2.24% | 4.375% | 11.8% | 9.187% | 12.43% | 55.58% | <LOD | 31.79% | 27.37% | 23.31% | 24.08% | <LOD | <LOD | 40.47% | 2.036% | <LOD | 70.76% | 26.49% | 0.3109%
6 [Col] | 2023-06-13 | NIST 679 | std | rep 1 | 12623 | 91028 | 218799 | <LOD | <LOD | 23414 | 1478 | 5520 | 215 | 108 | 1742 | 90490 | <LOD | 53 | 415 | 107 | 21 | 4 | 182 | 51 | 42 | 178 | 19 | <LOD | 44 | 66 | 60 | <LOD | 1927 | <LOD | 5 | 8 | <LOD | 40 | <LOD | <LOD | 704317
7 [Col] | 2023-06-13 | NIST 679 | std | rep 2 | 11756 | 92835 | 221562 | <LOD | <LOD | 23876 | 1416 | 5594 | 153 | 76 | 1631 | 90424 | 172 | 62 | 347 | 127 | 21 | 7 | 181 | 37 | 44 | 181 | 21 | <LOD | <LOD | 46 | 59 | 62 | 559 | <LOD | 8 | <LOD | <LOD | <LOD | 26 | <LOD | 702581
8 [Col] | 2023-06-13 | NIST 679 | std | rep 3 | <LOD | 87924 | 215456 | <LOD | <LOD | 23255 | 1435 | 5409 | 142 | 136 | 1705 | 90164 | <LOD | 67 | 406 | 136 | 24 | 5 | 189 | 53 | 47 | 162 | 17 | 9 | <LOD | <LOD | 64 | 94 | 637 | 35 | <LOD | 7 | <LOD | <LOD | 26 | 12 | 715017
9 [Col] | 2023-06-13 | NIST 679 | std | rep 4 | <LOD | 92004 | 222606 | <LOD | <LOD | 24295 | 1438 | 5548 | <LOD | 93 | 1578 | 93798 | 327 | 65 | 422 | 110 | 24 | 6 | 194 | 48 | 37 | 169 | 23 | 10 | 20 | 54 | 69 | 50 | 1238 | <LOD | <LOD | 11 | <LOD | <LOD | 32 | 6 | 700214
10 [Col] | 2023-06-13 | NIST 679 | std | rep 5 | 11348 | 93862 | 222884 | <LOD | <LOD | 24279 | 1634 | 5119 | 195 | 84 | 1731 | 91793 | 155 | 57 | 402 | 120 | 28 | 8 | 184 | 50 | 42 | 167 | 21 | 15 | <LOD | 63 | 54 | 56 | 480 | <LOD | 8 | 5 | <LOD | <LOD | 24 | <LOD | 699677
|  | NIST 679 | 2023-06-13 | average | 11909 | 91530.6 | 220261.4 | <LOD | <LOD | 23823.8 | 1480.2 | 5438 | 176.25 | 99.4 | 1677.4 | 91333.8 | 218 | 60.8 | 398.4 | 120 | 23.6 | 6 | 186 | 47.8 | 42.4 | 171.4 | 20.2 | 11.3333 | <LOD | 57.25 | 61.2 | 65.5 | 968.2 | <LOD | 7 | 7.75 | <LOD | <LOD | 27 | <LOD | 704361.2
|  | NIST 679 | 2023-06-13 | stdev | 651.1244 | 2270.5265 | 3134.3749 | <LOD | <LOD | 480.4765 | 88.8887 | 190.8939 | 34.4807 | 23.6601 | 70.4294 | 1515.8589 | 94.7787 | 5.7619 | 29.7708 | 11.9791 | 2.881 | 1.5811 | 5.4314 | 6.3008 | 3.6469 | 7.893 | 2.2804 | 3.2146 | <LOD | 9.0692 | 5.6303 | 19.6214 | 613.9151 | <LOD | 1.7321 | 2.5 | <LOD | <LOD | 3.4641 | <LOD | 6242.1578
|  | NIST 679 | 2023-06-13 | RSD | 5.467% | 2.481% | 1.423% | <LOD | <LOD | 2.017% | 6.005% | 3.51% | 19.56% | 23.8% | 4.199% | 1.66% | 43.48% | 9.477% | 7.473% | 9.983% | 12.21% | 26.35% | 2.92% | 13.18% | 8.601% | 4.605% | 11.29% | 28.36% | <LOD | 15.84% | 9.2% | 29.96% | 63.41% | <LOD | 24.74% | 32.26% | <LOD | <LOD | 12.83% | <LOD | 0.8862%
11 [Col] | 2023-06-13 | SARM 69 | std | rep 1 | 13795 | 76706 | 265469 | 331 | <LOD | 16968 | 14217 | 4503 | 163 | 224 | 1025 | 50699 | 96 | 42 | 321 | 66 | 16 | 2 | 72 | 96 | 28 | 292 | 10 | 12 | <LOD | <LOD | 59 | 66 | 867 | 25 | <LOD | 14 | <LOD | <LOD | 13 | <LOD | 709473
12 [Col] | 2023-06-13 | SARM 69 | std | rep 2 | 13517 | 78899 | 264042 | 353 | <LOD | 17398 | 14058 | 4596 | 153 | 172 | 961 | 51363 | <LOD | 72 | 338 | 62 | 7 | 6 | 70 | 105 | 29 | 315 | <LOD | 8 | 22 | 19 | 29 | 41 | 702 | 52 | <LOD | <LOD | <LOD | <LOD | <LOD | 5 | 708791
14 [Col] | 2023-06-13 | SARM 69 | std | rep 3 | 13063 | 77670 | 266629 | 319 | <LOD | 17509 | 13998 | 3632 | 76 | 218 | 999 | 51203 | 177 | 48 | 319 | 63 | 7 | 7 | 75 | 98 | 38 | 279 | 9 | 12 | <LOD | <LOD | 73 | 41 | 1032 | <LOD | <LOD | <LOD | 17 | 23 | <LOD | <LOD | 708665
15 [Col] | 2023-06-13 | SARM 69 | std | rep 4 | 13155 | 77484 | 265664 | 368 | <LOD | 17017 | 14467 | 4139 | 153 | 252 | 1009 | 51309 | 151 | 66 | 322 | 61 | 9 | 5 | 71 | 104 | 30 | 323 | 18 | 5 | <LOD | <LOD | <LOD | 91 | 927 | <LOD | <LOD | 10 | <LOD | <LOD | <LOD | 10 | 708875
16 [Col] | 2023-06-13 | SARM 69 | std | rep 5 | 13995 | 79465 | 266676 | 340 | <LOD | 17609 | 15093 | 5014 | 148 | 207 | 1008 | 52095 | 168 | 71 | 315 | 82 | 7 | 5 | 75 | 93 | 30 | 276 | 6 | 11 | <LOD | 34 | 48 | 127 | 597 | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | 5 | 703267
|  | SARM 69 | 2023-06-13 | average | 13505 | 78044.8 | 265696 | 342.2 | <LOD | 17300.2 | 14366.6 | 4376.8 | 138.6 | 214.6 | 1000.4 | 51333.8 | 148 | 59.8 | 323 | 66.8 | 9.2 | 5 | 72.6 | 99.2 | 31 | 297 | 10.75 | 9.6 | <LOD | <LOD | 52.25 | 73.2 | 825 | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | 6.6667 | 707814.2
|  | SARM 69 | 2023-06-13 | stdev | 400.6894 | 1117.361 | 1074.6113 | 19.0447 | <LOD | 291.1524 | 444.752 | 519.9218 | 35.4161 | 29.031 | 23.9332 | 500.2431 | 36.3043 | 13.8636 | 8.8034 | 8.7006 | 3.8987 | 1.8708 | 2.3022 | 5.1672 | 4 | 21.1542 | 5.1235 | 3.0496 | <LOD | <LOD | 18.5719 | 36.5267 | 174.7355 | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | 2.8868 | 2560.8563
|  | SARM 69 | 2023-06-13 | RSD | 2.967% | 1.432% | 0.4045% | 5.565% | <LOD | 1.683% | 3.096% | 11.88% | 25.55% | 13.53% | 2.392% | 0.9745% | 24.53% | 23.18% | 2.726% | 13.02% | 42.38% | 37.42% | 3.171% | 5.209% | 12.9% | 7.123% | 47.66% | 31.77% | <LOD | <LOD | 35.54% | 49.9% | 21.18% | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | 43.3% | 0.3618%
1 [Col] | 2023-06-14 | NIST 2711a | std | rep 1 | 14514 | 69621 | 260878 | 152 | 209 | 25806 | 22636 | 3500 | 151 | 57 | 634 | 29837 | 112 | 26 | 368 | 396 | 105 | 5 | 122 | 260 | 37 | 269 | 15 | 4 | <LOD | 75 | 31 | 75 | 813 | <LOD | <LOD | 12 | 1358 | <LOD | <LOD | <LOD | 728369
2 [Col] | 2023-06-14 | NIST 2711a | std | rep 2 | 13762 | 69277 | 261147 | 167 | 178 | 25040 | 22754 | <LOD | <LOD | <LOD | 670 | 29289 | 192 | 32 | 377 | 400 | 161 | 5 | 125 | 266 | 39 | 292 | 21 | 15 | <LOD | 57 | 76 | 62 | 2350 | 50 | <LOD | <LOD | 1359 | 17 | <LOD | <LOD | 729215
3 [Col] | 2023-06-14 | NIST 2711a | std | rep 3 | 15094 | 70012 | 260175 | 158 | 98 | 25572 | 23570 | 3349 | <LOD | <LOD | 637 | 29014 | 102 | <LOD | 389 | 395 | 138 | 4 | 117 | 244 | 31 | 271 | 12 | 8 | 11 | 73 | 26 | 68 | 1421 | 34 | <LOD | 7 | 1368 | <LOD | 18 | <LOD | 728223
4 [Col] | 2023-06-14 | NIST 2711a | std | rep 4 | 14423 | 68886 | 261413 | 124 | 116 | 24787 | 22654 | 3525 | 86 | 147 | 622 | 27924 | <LOD | 20 | 361 | 357 | 119 | 5 | 112 | 258 | 38 | 309 | 24 | 11 | <LOD | 76 | 48 | 116 | 915 | 42 | <LOD | 10 | 1296 | 22 | <LOD | 4 | 731107
5 [Col] | 2023-06-14 | NIST 2711a | std | rep 5 | 15557 | 67096 | 262799 | 230 | 186 | 25287 | 22628 | 3107 | <LOD | 138 | 660 | 28728 | 200 | 25 | 381 | 395 | 141 | 3 | 120 | 266 | 25 | 329 | 21 | 12 | 24 | 50 | 45 | 97 | 1052 | 18 | <LOD | 8 | 1356 | <LOD | 11 | <LOD | 728699
|  | NIST 2711a | 2023-06-14 | average | 14670 | 68978.4 | 261282.4 | 166.2 | 157.4 | 25298.4 | 22848.4 | 3370.25 | <LOD | 114 | 644.6 | 28958.4 | 151.5 | 25.75 | 375.2 | 388.6 | 132.8 | 4.4 | 119.2 | 258.8 | 34 | 294 | 18.6 | 10 | <LOD | 66.2 | 45.2 | 83.6 | 1310.2 | 36 | <LOD | 9.25 | 1347.4 | <LOD | <LOD | <LOD | 729122.6
|  | NIST 2711a | 2023-06-14 | stdev | 684.9004 | 1131.6352 | 965.0569 | 39.1178 | 47.8205 | 406.4927 | 406.54 | 191.9503 | <LOD | 49.5681 | 19.7687 | 708.4125 | 51.6495 | 4.9244 | 10.9636 | 17.7848 | 21.5221 | 0.89443 | 4.9699 | 9.0111 | 5.9161 | 25.5343 | 4.9295 | 4.1833 | <LOD | 11.9038 | 19.5371 | 22.4343 | 625.2037 | 13.6626 | <LOD | 2.2174 | 29.0998 | <LOD | <LOD | <LOD | 1172.9249
|  | NIST 2711a | 2023-06-14 | RSD | 4.669% | 1.641% | 0.3694% | 23.54% | 30.38% | 1.607% | 1.779% | 5.695% | <LOD | 43.48% | 3.067% | 2.446% | 34.09% | 19.12% | 2.922% | 4.577% | 16.21% | 20.33% | 4.169% | 3.482% | 17.4% | 8.685% | 26.5% | 41.83% | <LOD | 17.98% | 43.22% | 26.84% | 47.72% | 37.95% | <LOD | 23.97% | 2.16% | <LOD | <LOD | <LOD | 0.1609%
6 [Col] | 2023-06-14 | NIST 679 | std | rep 1 | 11288 | 93410 | 221930 | <LOD | <LOD | 24021 | 1455 | 5406 | 270 | 119 | 1458 | 93463 | <LOD | 52 | 373 | 112 | 21 | 3 | 183 | 47 | 45 | 171 | 22 | 18 | 28 | 25 | 56 | 73 | <LOD | 41 | <LOD | 16 | <LOD | <LOD | <LOD | <LOD | 700069
7 [Col] | 2023-06-14 | NIST 679 | std | rep 2 | <LOD | 91158 | 219610 | <LOD | <LOD | 24175 | 1481 | 5849 | 136 | 147 | 1546 | 93895 | 116 | 47 | 403 | 120 | 25 | 4 | 197 | 50 | 42 | 162 | 23 | 9 | 15 | <LOD | <LOD | 36 | <LOD | 34 | 4 | <LOD | <LOD | <LOD | 37 | <LOD | 704703
8 [Col] | 2023-06-14 | NIST 679 | std | rep 3 | 11661 | 94568 | 222879 | <LOD | <LOD | 24054 | 1374 | 4994 | 155 | <LOD | 1577 | 91039 | <LOD | 83 | 428 | 144 | 20 | 3 | 188 | 43 | 35 | 157 | 20 | 9 | <LOD | 18 | 70 | 87 | 767 | <LOD | 9 | 8 | 16 | <LOD | 23 | <LOD | 700031
9 [Col] | 2023-06-14 | NIST 679 | std | rep 4 | 11180 | 91140 | 216874 | <LOD | <LOD | 23287 | 1344 | 5497 | 261 | 112 | 1763 | 90819 | <LOD | 72 | 442 | 122 | 16 | 7 | 182 | 54 | 42 | 162 | 26 | 11 | 26 | 35 | 61 | 88 | 951 | 20 | <LOD | 19 | <LOD | <LOD | 13 | <LOD | 708327
10 [Col] | 2023-06-14 | NIST 679 | std | rep 5 | 11945 | 92074 | 221004 | <LOD | <LOD | 23542 | 1402 | 5109 | 186 | 126 | 1513 | 91839 | <LOD | 81 | 386 | 122 | 23 | <LOD | 197 | 48 | 38 | 162 | 27 | 16 | 74 | 47 | 71 | 66 | 1144 | 44 | <LOD | <LOD | 22 | <LOD | 13 | <LOD | 702200
|  | NIST 679 | 2023-06-14 | average | 11518.5 | 92470 | 220459.4 | <LOD | <LOD | 23815.8 | 1411.2 | 5371 | 201.6 | 126 | 1571.4 | 92211 | <LOD | 67 | 406.4 | 124 | 21 | 4.25 | 189.4 | 48.4 | 40.4 | 162.8 | 23.6 | 12.6 | 35.75 | 31.25 | 64.5 | 70 | 954 | 34.75 | <LOD | 14.3333 | <LOD | <LOD | 21.5 | <LOD | 703066
|  | NIST 679 | 2023-06-14 | stdev | 351.151 | 1494.3447 | 2339.3531 | <LOD | <LOD | 381.5976 | 56.5128 | 337.7714 | 61.0844 | 15.1217 | 115.8201 | 1401.1595 | <LOD | 16.5982 | 28.6234 | 11.9164 | 3.3912 | 1.893 | 7.3007 | 4.0373 | 3.9115 | 5.0695 | 2.881 | 4.1593 | 26.1327 | 12.6062 | 7.2342 | 21.1778 | 188.5179 | 10.6888 | <LOD | 5.6862 | <LOD | <LOD | 11.3578 | <LOD | 3510.7364
|  | NIST 679 | 2023-06-14 | RSD | 3.049% | 1.616% | 1.061% | <LOD | <LOD | 1.602% | 4.005% | 6.289% | 30.3% | 12% | 7.371% | 1.52% | <LOD | 24.77% | 7.043% | 9.61% | 16.15% | 44.54% | 3.855% | 8.342% | 9.682% | 3.114% | 12.21% | 33.01% | 73.1% | 40.34% | 11.22% | 30.25% | 19.76% | 30.76% | <LOD | 39.67% | <LOD | <LOD | 52.83% | <LOD | 0.4993%
11 [Col] | 2023-06-14 | SARM 69 | std | rep 1 | 12560 | 77115 | 262716 | 299 | <LOD | 17200 | 13598 | 4180 | 134 | 238 | 1031 | 51261 | 201 | 68 | 320 | 61 | 5 | 5 | 82 | 105 | 35 | 284 | 6 | 14 | 18 | 40 | 52 | 53 | 1816 | 33 | <LOD | <LOD | 21 | 20 | <LOD | <LOD | 712065
12 [Col] | 2023-06-14 | SARM 69 | std | rep 2 | 13010 | 76707 | 261540 | 226 | <LOD | 17259 | 13685 | 4271 | 96 | 146 | 957 | 50681 | 150 | 73 | 292 | 69 | 14 | 5 | 81 | 116 | 35 | 303 | 9 | 13 | <LOD | 14 | 37 | 58 | 1871 | <LOD | <LOD | <LOD | <LOD | 19 | <LOD | <LOD | 713691
13 [Col] | 2023-06-14 | SARM 69 | std | rep 3 | 13069 | 77628 | 263263 | 344 | <LOD | 17570 | 13832 | 4381 | 159 | 116 | 1026 | 51397 | 99 | 37 | 316 | 61 | 5 | 3 | 77 | 98 | 35 | 303 | 15 | 6 | 32 | 37 | 46 | 39 | 817 | 30 | <LOD | <LOD | <LOD | <LOD | 15 | <LOD | 711070
14 [Col] | 2023-06-14 | SARM 69 | std | rep 4 | 12900 | 77463 | 264651 | 412 | <LOD | 17490 | 14606 | 4927 | 155 | 335 | 960 | 51507 | 215 | 72 | 329 | 78 | 7 | 4 | 80 | 109 | 34 | 296 | <LOD | 6 | 29 | 18 | 85 | 104 | 829 | 46 | <LOD | <LOD | 18 | 20 | <LOD | <LOD | 708457
15 [Col] | 2023-06-14 | SARM 69 | std | rep 5 | 13411 | 77436 | 261143 | 287 | <LOD | 17471 | 14485 | 4143 | 102 | 213 | 862 | 50336 | 212 | 71 | 316 | 67 | 9 | 4 | 77 | 100 | 28 | 246 | 12 | 5 | <LOD | 43 | <LOD | 36 | 1051 | <LOD | 4 | 6 | <LOD | 36 | <LOD | <LOD | 713671
|  | SARM 69 | 2023-06-14 | average | 12990 | 77269.8 | 262662.6 | 313.6 | <LOD | 17398 | 14041.2 | 4380.4 | 129.2 | 209.6 | 967.2 | 51036.4 | 175.4 | 64.2 | 314.6 | 67.2 | 8 | 4.2 | 79.4 | 105.6 | 33.4 | 286.4 | 10.5 | 8.8 | 26.3333 | 30.4 | 55 | 58 | 1276.8 | 36.3333 | <LOD | <LOD | <LOD | 23.75 | <LOD | <LOD | 711790.8
|  | SARM 69 | 2023-06-14 | stdev | 306.9536 | 365.3638 | 1404.144 | 69.2842 | <LOD | 159.6105 | 469.8475 | 319.0843 | 29.2353 | 85.658 | 68.4668 | 504.6809 | 50.1129 | 15.3199 | 13.704 | 7.0143 | 3.7417 | 0.83666 | 2.3022 | 7.2319 | 3.0496 | 23.881 | 3.873 | 4.3243 | 7.3711 | 13.3903 | 20.9284 | 27.3222 | 526.0078 | 8.5049 | <LOD | <LOD | <LOD | 8.1803 | <LOD | <LOD | 2171.1099
|  | SARM 69 | 2023-06-14 | RSD | 2.363% | 0.4728% | 0.5346% | 22.09% | <LOD | 0.9174% | 3.346% | 7.284% | 22.63% | 40.87% | 7.079% | 0.9889% | 28.57% | 23.86% | 4.356% | 10.44% | 46.77% | 19.92% | 2.899% | 6.848% | 9.131% | 8.338% | 36.89% | 49.14% | 27.99% | 44.05% | 38.05% | 47.11% | 41.2% | 23.41% | <LOD | <LOD | <LOD | 34.44% | <LOD | <LOD | 0.305%
1 [Col] | 2023-06-15 | NIST 2711a | std | rep 1 | 14406 | 70601 | 266886 | 194 | 101 | 25791 | 23140 | 3702 | 79 | 118 | 697 | 29017 | 66 | 20 | 271 | 395 | 161 | 5 | 121 | 258 | 28 | 323 | 18 | 6 | 16 | 87 | 51 | 60 | 737 | 38 | <LOD | <LOD | 1361 | <LOD | <LOD | <LOD | 722563
2 [Col] | 2023-06-15 | NIST 2711a | std | rep 2 | 14507 | 69332 | 262826 | 204 | 170 | 25689 | 23268 | 3291 | 127 | 115 | 794 | 29907 | 95 | 28 | 316 | 381 | 143 | 4 | 131 | 258 | 30 | 342 | 22 | 16 | 20 | 97 | 74 | 102 | 954 | 25 | 7 | 9 | 1372 | <LOD | 8 | <LOD | 726077
3 [Col] | 2023-06-15 | NIST 2711a | std | rep 3 | 15879 | 69400 | 264460 | 161 | 130 | 24858 | 23022 | 3356 | <LOD | 73 | 697 | 28784 | 66 | 32 | 261 | 377 | 100 | 4 | 116 | 266 | 36 | 306 | 26 | 12 | <LOD | 73 | 56 | 64 | 1780 | 34 | <LOD | 9 | 1338 | <LOD | 10 | <LOD | 724445
4 [Col] | 2023-06-15 | NIST 2711a | std | rep 4 | 15055 | 69317 | 263373 | 123 | 128 | 25377 | 22040 | <LOD | <LOD | 95 | 779 | 28625 | <LOD | 37 | 297 | 370 | 98 | 2 | 122 | 261 | 34 | 264 | 18 | 11 | 25 | 90 | 47 | 91 | 1027 | 21 | 4 | 6 | 1349 | <LOD | 9 | <LOD | 728047
5 [Col] | 2023-06-15 | NIST 2711a | std | rep 5 | 14797 | 69450 | 264869 | 221 | 127 | 26161 | 23097 | 3110 | 85 | 57 | 632 | 29923 | <LOD | 34 | 248 | 411 | 148 | 6 | 117 | 268 | 32 | 311 | 9 | 6 | 24 | 77 | 51 | 91 | 1315 | 25 | 6 | <LOD | 1373 | <LOD | 20 | <LOD | 723916
|  | NIST 2711a | 2023-06-15 | average | 14928.8 | 69620 | 264482.8 | 180.6 | 131.2 | 25575.2 | 22913.4 | 3364.75 | 97 | 91.6 | 719.8 | 29251.2 | 75.6667 | 30.2 | 278.6 | 386.8 | 130 | 4.2 | 121.4 | 262.2 | 32 | 309.2 | 18.6 | 10.2 | 21.25 | 84.8 | 55.8 | 81.6 | 1162.6 | 28.6 | 5.6667 | 8 | 1358.6 | <LOD | 11.75 | <LOD | 725009.6
|  | NIST 2711a | 2023-06-15 | stdev | 588.9424 | 551.0023 | 1573.347 | 38.9269 | 24.7326 | 488.9634 | 496.3414 | 247.7571 | 26.1534 | 26.4537 | 66.6311 | 621.8201 | 16.7432 | 6.5727 | 27.5735 | 16.3156 | 29.0603 | 1.4832 | 5.9414 | 4.6043 | 3.1623 | 28.8219 | 6.3087 | 4.2661 | 4.113 | 9.757 | 10.663 | 18.5014 | 402.3199 | 7.0922 | 1.5275 | 1.7321 | 15.0765 | <LOD | 5.5603 | <LOD | 2113.3934
|  | NIST 2711a | 2023-06-15 | RSD | 3.945% | 0.7914% | 0.5949% | 21.55% | 18.85% | 1.912% | 2.166% | 7.363% | 26.96% | 28.88% | 9.257% | 2.126% | 22.13% | 21.76% | 9.897% | 4.218% | 22.35% | 35.32% | 4.894% | 1.756% | 9.882% | 9.321% | 33.92% | 41.82% | 19.36% | 11.51% | 19.11% | 22.67% | 34.61% | 24.8% | 26.96% | 21.65% | 1.11% | <LOD | 47.32% | <LOD | 0.2915%
6 [Col] | 2023-06-15 | NIST 679 | std | rep 1 | 11574 | 90355 | 216641 | <LOD | <LOD | 23352 | 1467 | 5431 | 81 | 100 | 1621 | 89805 | 326 | 51 | 219 | 114 | 23 | 4 | 197 | 55 | 39 | 163 | 21 | 15 | 24 | 27 | 61 | 138 | 961 | 32 | <LOD | <LOD | <LOD | 20 | <LOD | <LOD | 709885
7 [Col] | 2023-06-15 | NIST 679 | std | rep 2 | <LOD | 89292 | 216964 | <LOD | <LOD | 23585 | 1423 | 5797 | 167 | 131 | 1823 | 92308 | <LOD | 79 | 260 | 113 | 24 | 6 | 196 | 55 | 39 | 155 | 21 | 23 | <LOD | 44 | 74 | 45 | 608 | 46 | <LOD | <LOD | <LOD | <LOD | 24 | <LOD | 709846
8 [Col] | 2023-06-15 | NIST 679 | std | rep 3 | <LOD | 94512 | 223310 | <LOD | <LOD | 24433 | 1459 | 5536 | 155 | 126 | 1674 | 94985 | 196 | 63 | 231 | 120 | 24 | 2 | 206 | 57 | 47 | 179 | 22 | 20 | 28 | 29 | <LOD | 78 | 748 | 48 | <LOD | <LOD | <LOD | <LOD | 30 | <LOD | 696835
9 [Col] | 2023-06-15 | NIST 679 | std | rep 4 | 11058 | 92185 | 221121 | <LOD | <LOD | 24018 | 1449 | 5470 | 283 | 50 | 1484 | 90475 | 177 | 49 | 244 | 121 | 19 | 3 | 189 | 49 | 48 | 157 | 30 | 9 | 15 | 34 | 105 | 114 | 476 | 41 | <LOD | <LOD | 25 | <LOD | 16 | 8 | 704375
10 [Col] | 2023-06-15 | NIST 679 | std | rep 5 | 12225 | 91954 | 221134 | <LOD | <LOD | 23924 | 1403 | 5982 | <LOD | 82 | 1569 | 92175 | 162 | 74 | 257 | 132 | 23 | 7 | 187 | 56 | 38 | 165 | 35 | 24 | 21 | 41 | 50 | 112 | 827 | <LOD | <LOD | 13 | <LOD | <LOD | 9 | <LOD | 700865
|  | NIST 679 | 2023-06-15 | average | 11619 | 91659.6 | 219834 | <LOD | <LOD | 23862.4 | 1440.2 | 5643.2 | 171.5 | 97.8 | 1634.2 | 91949.6 | 215.25 | 63.2 | 242.2 | 120 | 22.6 | 4.4 | 195 | 54.4 | 42.2 | 163.8 | 25.8 | 18.2 | 22 | 35 | 72.5 | 97.4 | 724 | 41.75 | <LOD | <LOD | <LOD | <LOD | 19.75 | <LOD | 704361.2
|  | NIST 679 | 2023-06-15 | stdev | 584.8 | 1987.864 | 2909.5143 | <LOD | <LOD | 415.6468 | 26.5932 | 237.1491 | 83.4965 | 33.3047 | 126.7032 | 2010.3895 | 75.1327 | 13.3866 | 17.3407 | 7.5829 | 2.0736 | 2.0736 | 7.5166 | 3.1305 | 4.8683 | 9.4446 | 6.3797 | 6.2209 | 5.4772 | 7.3824 | 23.7837 | 36.2602 | 188.6226 | 7.1356 | <LOD | <LOD | <LOD | <LOD | 9.1788 | <LOD | 5689.0836
|  | NIST 679 | 2023-06-15 | RSD | 5.033% | 2.169% | 1.324% | <LOD | <LOD | 1.742% | 1.846% | 4.202% | 48.69% | 34.05% | 7.753% | 2.186% | 34.9% | 21.18% | 7.16% | 6.319% | 9.175% | 47.13% | 3.855% | 5.755% | 11.54% | 5.766% | 24.73% | 34.18% | 24.9% | 21.09% | 32.81% | 37.23% | 26.05% | 17.09% | <LOD | <LOD | <LOD | <LOD | 46.47% | <LOD | 0.8077%
11 [Col] | 2023-06-15 | SARM 69 | std | rep 1 | 13995 | 78039 | 266307 | 239 | <LOD | 17535 | 14623 | 4409 | 270 | 200 | 961 | 51043 | 93 | 65 | 177 | 62 | 7 | 7 | 73 | 107 | 35 | 361 | 10 | 11 | 30 | 37 | 43 | 36 | <LOD | <LOD | 6 | <LOD | 17 | 23 | <LOD | <LOD | 707571
12 [Col] | 2023-06-15 | SARM 69 | std | rep 2 | 13442 | 78961 | 266243 | 320 | <LOD | 17688 | 13807 | 4567 | 124 | 193 | 1006 | 50732 | 63 | 51 | 179 | 72 | 6 | 4 | 76 | 106 | 35 | 307 | 11 | 15 | <LOD | 24 | 54 | 71 | 529 | 26 | <LOD | 10 | <LOD | <LOD | 13 | 4 | 707672
13 [Col] | 2023-06-15 | SARM 69 | std | rep 3 | 13222 | 78058 | 265269 | 252 | <LOD | 17463 | 14277 | 4467 | 78 | 261 | 1079 | 50697 | 193 | 44 | 210 | 57 | 10 | 3 | 80 | 113 | 37 | 262 | 8 | 7 | 13 | 52 | 49 | 35 | 1532 | 36 | <LOD | 8 | <LOD | <LOD | 16 | <LOD | 708383
14 [Col] | 2023-06-15 | SARM 69 | std | rep 4 | 12711 | 79497 | 269757 | 331 | <LOD | 17918 | 14094 | 4418 | 167 | 241 | 943 | 51999 | 136 | 49 | 200 | 79 | 16 | 4 | 76 | 106 | 29 | 265 | <LOD | <LOD | <LOD | 15 | 54 | 89 | 617 | <LOD | <LOD | 5 | <LOD | <LOD | <LOD | <LOD | 703313
15 [Col] | 2023-06-15 | SARM 69 | std | rep 5 | 13255 | 77463 | 264568 | 271 | <LOD | 17723 | 14086 | 4653 | 124 | 200 | 1019 | 52020 | 238 | 77 | 207 | 74 | 4 | 3 | 85 | 106 | 33 | 276 | 8 | 11 | <LOD | <LOD | 54 | 75 | 1028 | 17 | <LOD | 10 | 16 | <LOD | 29 | <LOD | 708366
|  | SARM 69 | 2023-06-15 | average | 13325 | 78403.6 | 266428.8 | 282.6 | <LOD | 17665.4 | 14177.4 | 4502.8 | 152.6 | 219 | 1001.6 | 51298.2 | 144.6 | 57.2 | 194.6 | 68.8 | 8.6 | 4.2 | 78 | 107.6 | 33.8 | 294.2 | 9.25 | 11 | <LOD | 32 | 50.8 | 61.2 | 926.5 | 26.3333 | <LOD | 8.25 | <LOD | <LOD | 19.3333 | <LOD | 707061
|  | SARM 69 | 2023-06-15 | stdev | 462.3565 | 812.9144 | 1995.8698 | 40.9671 | <LOD | 177.2041 | 300.4102 | 104.8246 | 72.786 | 30.1911 | 53.3742 | 663.1868 | 71.5213 | 13.5351 | 15.5981 | 9.0388 | 4.669 | 1.6432 | 4.6368 | 3.0496 | 3.0332 | 41.3727 | 1.5 | 3.266 | <LOD | 16.1038 | 4.8683 | 24.3967 | 458.5226 | 9.5044 | <LOD | 2.3629 | <LOD | <LOD | 8.5049 | <LOD | 2129.0628
|  | SARM 69 | 2023-06-15 | RSD | 3.47% | 1.037% | 0.7491% | 14.5% | <LOD | 1.003% | 2.119% | 2.328% | 47.7% | 13.79% | 5.329% | 1.293% | 49.46% | 23.66% | 8.015% | 13.14% | 54.29% | 39.12% | 5.945% | 2.834% | 8.974% | 14.06% | 16.22% | 29.69% | <LOD | 50.32% | 9.583% | 39.86% | 49.49% | 36.09% | <LOD | 28.64% | <LOD | <LOD | 43.99% | <LOD | 0.3011%
1 [Col] | 2023-06-16 | NIST 2711a | std | rep 1 | 14384 | 66551 | 259347 | 82 | 173 | 25044 | 22069 | 3362 | 137 | 84 | 629 | 28466 | 77 | 27 | 303 | 383 | 126 | 4 | 122 | 253 | 31 | 285 | 18 | 13 | <LOD | 81 | 36 | 79 | 1296 | <LOD | <LOD | 8 | 1355 | <LOD | 13 | <LOD | 734332
2 [Col] | 2023-06-16 | NIST 2711a | std | rep 2 | 15319 | 68885 | 261955 | 108 | 109 | 25497 | 23231 | 3742 | <LOD | 84 | 706 | 29432 | 75 | 27 | 327 | 411 | 146 | <LOD | 127 | 263 | 35 | 347 | 26 | <LOD | 33 | 77 | 36 | 85 | 1363 | 43 | <LOD | 13 | 1380 | <LOD | <LOD | 8 | 726419
3 [Col] | 2023-06-16 | NIST 2711a | std | rep 3 | 14470 | 68749 | 258949 | 94 | 100 | 25251 | 22103 | 3313 | 139 | 100 | 673 | 28061 | <LOD | 29 | 251 | 372 | 89 | 6 | 125 | 255 | 31 | 283 | 22 | 7 | 30 | 86 | 50 | 101 | <LOD | <LOD | 5 | 7 | 1334 | 24 | <LOD | <LOD | 734604
4 [Col] | 2023-06-16 | NIST 2711a | std | rep 4 | 15494 | 69119 | 263320 | 175 | 229 | 25460 | 23137 | 4278 | 196 | 104 | 668 | 30162 | <LOD | 24 | 320 | 395 | 176 | <LOD | 123 | 268 | 33 | 337 | 24 | 10 | <LOD | 106 | 59 | 80 | 614 | 41 | <LOD | <LOD | 1356 | <LOD | 13 | <LOD | 724037
5 [Col] | 2023-06-16 | NIST 2711a | std | rep 5 | 14739 | 69461 | 261636 | 166 | 117 | 25422 | 22445 | 3150 | <LOD | 69 | 565 | 29808 | 75 | 32 | 263 | 391 | 173 | 3 | 111 | 268 | 32 | 307 | 18 | 7 | 22 | 97 | 50 | 117 | 1332 | <LOD | <LOD | 16 | 1301 | <LOD | 20 | <LOD | 728028
|  | NIST 2711a | 2023-06-16 | average | 14881.2 | 68553 | 261041.4 | 125 | 145.6 | 25334.8 | 22597 | 3569 | 157.3333 | 88.2 | 648.2 | 29185.8 | 75.6667 | 27.8 | 292.8 | 390.4 | 142 | 4.3333 | 121.6 | 261.4 | 32.4 | 311.8 | 21.6 | 9.25 | 28.3333 | 89.4 | 46.2 | 92.4 | 1151.25 | <LOD | <LOD | 11 | 1345.2 | <LOD | 15.3333 | <LOD | 729484
|  | NIST 2711a | 2023-06-16 | stdev | 500.9258 | 1151.2454 | 1845.8982 | 42.6615 | 54.6425 | 187.8662 | 556.6597 | 451.8174 | 33.5012 | 14.0784 | 53.9416 | 892.1884 | 1.1547 | 2.9496 | 34.0911 | 14.4845 | 36.0486 | 1.5275 | 6.229 | 7.0922 | 1.6733 | 29.3462 | 3.5777 | 2.8723 | 5.6862 | 11.929 | 10.01 | 16.334 | 359.2115 | <LOD | <LOD | 4.2426 | 29.5922 | <LOD | 4.0415 | <LOD | 4767.1132
|  | NIST 2711a | 2023-06-16 | RSD | 3.366% | 1.679% | 0.7071% | 34.13% | 37.53% | 0.7415% | 2.463% | 12.66% | 21.29% | 15.96% | 8.322% | 3.057% | 1.526% | 10.61% | 11.64% | 3.71% | 25.39% | 35.25% | 5.123% | 2.713% | 5.165% | 9.412% | 16.56% | 31.05% | 20.07% | 13.34% | 21.67% | 17.68% | 31.2% | <LOD | <LOD | 38.57% | 2.2% | <LOD | 26.36% | <LOD | 0.6535%
6 [Col] | 2023-06-16 | NIST 679 | std | rep 1 | 12730 | 94004 | 223819 | <LOD | <LOD | 24236 | 1413 | 6303 | 345 | 45 | 1706 | 92541 | 184 | 66 | 264 | 142 | 23 | 9 | 193 | 53 | 38 | 177 | 24 | 16 | <LOD | 48 | 37 | 35 | 577 | 54 | <LOD | <LOD | <LOD | 30 | <LOD | 6 | 695096
7 [Col] | 2023-06-16 | NIST 679 | std | rep 2 | 10938 | 92600 | 221779 | <LOD | <LOD | 24070 | 1463 | 5217 | 322 | 169 | 1374 | 89439 | 152 | 68 | 229 | 167 | 26 | 4 | 184 | 50 | 40 | 154 | 25 | 8 | 18 | 42 | 82 | 36 | <LOD | <LOD | 4 | <LOD | 19 | <LOD | 12 | 6 | 705402
8 [Col] | 2023-06-16 | NIST 679 | std | rep 3 | <LOD | 95182 | 221733 | <LOD | <LOD | 24061 | 1476 | 5487 | 105 | 124 | 1617 | 93442 | 145 | 42 | 237 | 127 | 15 | 7 | 199 | 62 | 34 | 169 | 24 | 17 | 24 | 17 | 29 | 50 | 1247 | 20 | <LOD | 8 | 21 | <LOD | <LOD | <LOD | 699339
9 [Col] | 2023-06-16 | NIST 679 | std | rep 4 | 13102 | 91256 | 221472 | <LOD | <LOD | 24136 | 1492 | 5958 | 137 | 123 | 1711 | 93493 | <LOD | 81 | 227 | 110 | 28 | 4 | 195 | 51 | 38 | 169 | 22 | 5 | 22 | 37 | 87 | 57 | 539 | 16 | <LOD | <LOD | <LOD | <LOD | 19 | <LOD | 698803
10 [Col] | 2023-06-16 | NIST 679 | std | rep 5 | 11808 | 92052 | 220513 | <LOD | <LOD | 24044 | 1473 | 5214 | 148 | 41 | 1619 | 93357 | <LOD | 45 | 207 | 124 | 24 | 3 | 185 | 40 | 30 | 160 | 24 | 10 | <LOD | 57 | 71 | 93 | 666 | 18 | <LOD | <LOD | <LOD | <LOD | 17 | 6 | 701673
|  | NIST 679 | 2023-06-16 | average | 12144.5 | 93018.8 | 221863.2 | <LOD | <LOD | 24109.4 | 1463.4 | 5635.8 | 211.4 | 100.4 | 1605.4 | 92454.4 | 160.3333 | 60.4 | 232.8 | 134 | 23.2 | 5.4 | 191.2 | 51.2 | 36 | 165.8 | 23.8 | 11.2 | 21.3333 | 40.2 | 61.2 | 54.2 | 757.25 | 27 | <LOD | <LOD | <LOD | <LOD | 16 | 6 | 700062.6
|  | NIST 679 | 2023-06-16 | stdev | 970.9945 | 1570.6378 | 1206.7801 | <LOD | <LOD | 78.9101 | 30.0383 | 480.7335 | 112.8685 | 55.6129 | 137.0558 | 1729.8025 | 20.7926 | 16.5015 | 20.6446 | 21.6679 | 4.9699 | 2.51 | 6.4962 | 7.8549 | 4 | 8.9275 | 1.0954 | 5.1672 | 3.0551 | 14.9566 | 26.5368 | 23.6157 | 330.8095 | 18.0739 | <LOD | <LOD | <LOD | <LOD | 3.6056 | 0 | 3803.9415
|  | NIST 679 | 2023-06-16 | RSD | 7.995% | 1.689% | 0.5439% | <LOD | <LOD | 0.3273% | 2.053% | 8.53% | 53.39% | 55.39% | 8.537% | 1.871% | 12.97% | 27.32% | 8.868% | 16.17% | 21.42% | 46.48% | 3.398% | 15.34% | 11.11% | 5.384% | 4.603% | 46.14% | 14.32% | 37.21% | 43.36% | 43.57% | 43.69% | 66.94% | <LOD | <LOD | <LOD | <LOD | 22.53% | 0% | 0.5434%
11 [Col] | 2023-06-16 | SARM 69 | std | rep 1 | 12628 | 78743 | 267030 | 302 | <LOD | 17250 | 14035 | 4809 | 99 | 228 | 856 | 50693 | <LOD | 55 | 194 | 59 | 13 | 3 | 78 | 99 | 32 | 259 | 9 | <LOD | 44 | 61 | 35 | 27 | 570 | 33 | <LOD | 4 | <LOD | <LOD | <LOD | <LOD | 708200
12 [Col] | 2023-06-16 | SARM 69 | std | rep 2 | 13066 | 79366 | 269789 | 392 | <LOD | 17653 | 14746 | 4724 | 108 | 184 | 964 | 52161 | 255 | 61 | 221 | 73 | 6 | 7 | 75 | 111 | 34 | 301 | 7 | 7 | 14 | 35 | 59 | 69 | 402 | <LOD | <LOD | <LOD | 21 | <LOD | 22 | 4 | 702234
13 [Col] | 2023-06-16 | SARM 69 | std | rep 3 | 11538 | 78915 | 266501 | 320 | <LOD | 17407 | 14066 | 4503 | 249 | 293 | 1000 | 51206 | 71 | 73 | 216 | 73 | 5 | 2 | 79 | 109 | 35 | 299 | 15 | 11 | 12 | 25 | 36 | 90 | 1128 | <LOD | 8 | 8 | 17 | 42 | <LOD | <LOD | 708378
14 [Col] | 2023-06-16 | SARM 69 | std | rep 4 | 13429 | 78984 | 266633 | 275 | <LOD | 17858 | 14356 | 4909 | 178 | 198 | 1018 | 51021 | 273 | 56 | 235 | 78 | 5 | 3 | 76 | 104 | 29 | 332 | 8 | 14 | <LOD | <LOD | 64 | 67 | <LOD | 25 | <LOD | 9 | <LOD | <LOD | 8 | <LOD | 706452
15 [Col] | 2023-06-16 | SARM 69 | std | rep 5 | 12648 | 78607 | 269833 | 296 | <LOD | 17704 | 14409 | 4361 | 155 | 161 | 992 | 50157 | 186 | 53 | 231 | 79 | 4 | 8 | 77 | 107 | 27 | 309 | 8 | <LOD | 17 | 27 | 57 | 79 | 881 | <LOD | <LOD | <LOD | 17 | <LOD | 11 | <LOD | 705505
|  | SARM 69 | 2023-06-16 | average | 12661.8 | 78923 | 267957.2 | 317 | <LOD | 17574.4 | 14322.4 | 4661.2 | 157.8 | 212.8 | 966 | 51047.6 | 196.25 | 59.6 | 219.4 | 72.4 | 6.6 | 4.6 | 77 | 106 | 31.4 | 300 | 9.4 | 10.6667 | 21.75 | 37 | 50.2 | 66.4 | 745.25 | <LOD | <LOD | 7 | 18.3333 | <LOD | 13.6667 | <LOD | 706153.8
|  | SARM 69 | 2023-06-16 | stdev | 709.9706 | 288.223 | 1703.5141 | 44.8999 | <LOD | 243.2392 | 290.0022 | 224.9115 | 60.5615 | 50.9774 | 64.4981 | 738.9802 | 91.5328 | 8.0498 | 16.1028 | 7.9875 | 3.6469 | 2.7019 | 1.5811 | 4.6904 | 3.3615 | 26.4008 | 3.2094 | 3.5119 | 14.975 | 16.5731 | 13.6638 | 23.8495 | 323.2433 | <LOD | <LOD | 2.6458 | 2.3094 | <LOD | 7.3711 | <LOD | 2500.4348
|  | SARM 69 | 2023-06-16 | RSD | 5.607% | 0.3652% | 0.6357% | 14.16% | <LOD | 1.384% | 2.025% | 4.825% | 38.38% | 23.96% | 6.677% | 1.448% | 46.64% | 13.51% | 7.339% | 11.03% | 55.26% | 58.74% | 2.053% | 4.425% | 10.71% | 8.8% | 34.14% | 32.92% | 68.85% | 44.79% | 27.22% | 35.92% | 43.37% | <LOD | <LOD | 37.8% | 12.6% | <LOD | 53.93% | <LOD | 0.3541%
2-RP1 [Col] | 2023-07-18 | NIST 2711a | std | rep 1 | 12475.8 | 65610.4 | 256441.7 | <LOD | <LOD | 23905.4 | 21743.3 | <LOD | 113.9 | 34.4 | 614.2 | 27652 | 161.9 | 26.4 | 172.2 | 364.3 | 118.4 | 2 | 114.2 | 257.1 | 36.9 | 267.9 | 20.6 | 10.8 | 22.5 | 93.7 | 77.3 | 117.2 | 1038.9 | <LOD | <LOD | 7.2 | 1274.1 | <LOD | 11.3 | <LOD | 743061.1
3 [Col] | 2023-07-18 | NIST 2711a | std | rep 2 | <LOD | 63834.3 | 247201 | <LOD | <LOD | 24005.5 | 21803.1 | <LOD | <LOD | 105.7 | 630.4 | 27469.6 | 159.7 | 21.1 | 218.8 | 391.1 | 147.3 | <LOD | 125.3 | 252.4 | 28.5 | 325.1 | 16.6 | 8.6 | 37.6 | 82.7 | 34.9 | 77.4 | 1228 | <LOD | <LOD | <LOD | 1316.6 | 32.6 | <LOD | <LOD | 755655.9
4 [Col] | 2023-07-18 | NIST 2711a | std | rep 3 | 11619.1 | 66900.5 | 261294.4 | <LOD | <LOD | 25475.5 | 22424.2 | 3200.3 | 73.1 | 92.7 | 555.7 | 28425.1 | <LOD | <LOD | 202.6 | 399.1 | 107.7 | 3.5 | 113.2 | 253.3 | 30.6 | 277.5 | 25 | 9.1 | 13.7 | 61.3 | 31.8 | 93.2 | 847 | <LOD | <LOD | 4.4 | 1335.2 | 47.6 | <LOD | 4.1 | 736244.1
5 [Col] | 2023-07-18 | NIST 2711a | std | rep 4 | <LOD | 64800.2 | 250249 | <LOD | <LOD | 23505.5 | 22132.9 | 3275.4 | 258.3 | <LOD | 741.7 | 28430.5 | 151.5 | 24.6 | 196.6 | 356 | 154.3 | 2.9 | 118.4 | 257.6 | 35.6 | 292.9 | 21.7 | 20.6 | 12.9 | 86 | 29.1 | 114.6 | 1539.7 | 16.4 | <LOD | 14.5 | 1302.8 | <LOD | <LOD | <LOD | 750336.2
6 [Col] | 2023-07-18 | NIST 2711a | std | rep 5 | 11692.3 | 68982.8 | 265943.1 | <LOD | <LOD | 25462.1 | 22701.2 | 3740.4 | <LOD | <LOD | 688 | 28966.6 | 123.5 | 23 | 193.7 | 385.7 | 142.2 | 3.3 | 111.4 | 261.7 | 29.3 | 319.5 | 18.3 | 16.2 | <LOD | 86.9 | 59.6 | 108.1 | 878.3 | 41.8 | <LOD | 7.1 | 1285.1 | <LOD | 6.8 | 4.2 | 728763
|  | NIST 2711a | 2023-07-18 | average | 11929.0667 | 66025.64 | 256225.84 | <LOD | <LOD | 24470.8 | 22160.94 | 3405.3667 | 148.4333 | 77.6 | 646 | 28188.76 | 149.15 | 23.775 | 196.78 | 379.24 | 133.98 | 2.925 | 116.5 | 256.42 | 32.18 | 296.58 | 20.44 | 13.06 | 21.675 | 82.12 | 46.54 | 102.1 | 1106.38 | <LOD | <LOD | 8.3 | 1302.76 | <LOD | <LOD | <LOD | 742812.06
|  | NIST 2711a | 2023-07-18 | stdev | 474.8974 | 1999.1419 | 7702.7981 | <LOD | <LOD | 930.0649 | 407.5677 | 292.5671 | 97.3097 | 37.9728 | 71.2937 | 617.38 | 17.6759 | 2.2603 | 16.8259 | 18.3038 | 19.9454 | 0.66521 | 5.5507 | 3.7292 | 3.818 | 25.1935 | 3.23 | 5.1815 | 11.473 | 12.3066 | 21.0574 | 16.655 | 285.6614 | <LOD | <LOD | 4.3321 | 24.3781 | <LOD | <LOD | <LOD | 10748.8734
|  | NIST 2711a | 2023-07-18 | RSD | 3.981% | 3.028% | 3.006% | <LOD | <LOD | 3.801% | 1.839% | 8.591% | 65.56% | 48.93% | 11.04% | 2.19% | 11.85% | 9.507% | 8.551% | 4.826% | 14.89% | 22.74% | 4.765% | 1.454% | 11.86% | 8.495% | 15.8% | 39.67% | 52.93% | 14.99% | 45.25% | 16.31% | 25.82% | <LOD | <LOD | 52.19% | 1.871% | <LOD | <LOD | <LOD | 1.447%
7 [Col] | 2023-07-18 | NIST 679 | std | rep 1 | <LOD | 85224.5 | 214541.3 | <LOD | <LOD | 23038.6 | 1523.1 | 5225 | 111.3 | 91.9 | 1529.5 | 88109.3 | 215.2 | 49.5 | 106.2 | 142.1 | 17 | 4.4 | 187.9 | 49.5 | 44.6 | 169.3 | 18.8 | 9.4 | 45.3 | 39.9 | 59.1 | 80.2 | 1139.4 | 17.7 | <LOD | <LOD | 18.5 | <LOD | 28.1 | <LOD | 720117.7
8 [Col] | 2023-07-18 | NIST 679 | std | rep 2 | <LOD | 84741.8 | 211765.6 | <LOD | <LOD | 22495.5 | 1526.7 | 5254.3 | <LOD | 52 | 1522 | 89422.6 | 248.9 | 50.3 | 118.9 | 112.2 | 13.6 | 4.7 | 188.9 | 40.6 | 30.7 | 161.1 | 22.7 | 7.7 | 15.1 | 65.2 | 63.9 | 84 | 618.5 | 15.8 | <LOD | <LOD | 20.4 | <LOD | 27.3 | <LOD | 722720.5
9 [Col] | 2023-07-18 | NIST 679 | std | rep 3 | <LOD | 89634.3 | 219125.9 | <LOD | <LOD | 23663.7 | 1498.9 | 4804.2 | <LOD | 50.8 | 1587.5 | 91516.9 | 253.6 | 24.7 | 99.5 | 121.4 | 25.7 | <LOD | 194.8 | 49.6 | 43.7 | 168.9 | 35 | 13.8 | 29.4 | 69.7 | 27.7 | 45.7 | 1315.5 | 39.7 | <LOD | <LOD | <LOD | <LOD | 43.7 | <LOD | 709055.9
10 [Col] | 2023-07-18 | NIST 679 | std | rep 4 | <LOD | 92373.8 | 219717.3 | <LOD | <LOD | 23560.6 | 1447.4 | 5343.3 | 73 | 95.7 | 1740.5 | 88720.2 | 261.3 | 27.7 | 104.2 | 115.2 | 28.4 | 6.1 | 192.7 | 44.6 | 39 | 150.5 | 25.8 | 21.3 | 21.8 | 24.8 | 74.2 | 45.2 | 973.9 | 42.7 | <LOD | <LOD | <LOD | 21.2 | 8.8 | <LOD | 708796
11 [Col] | 2023-07-18 | NIST 679 | std | rep 5 | <LOD | 73652.7 | 183996 | <LOD | <LOD | 16881.1 | <LOD | 3561.2 | 106.2 | 56.4 | 995 | 56927.3 | 159.4 | 16.9 | 72.6 | 76 | 16.9 | 2.9 | 108.1 | <LOD | 27.2 | 87.3 | 11.6 | 13.9 | <LOD | 13 | 21.3 | 65.2 | 506 | 17.3 | <LOD | <LOD | <LOD | 22.3 | <LOD | <LOD | 796111
|  | NIST 679 | 2023-07-18 | average | <LOD | 85125.42 | 209829.22 | <LOD | <LOD | 21927.9 | 1499.025 | 4837.6 | 96.8333 | 69.36 | 1474.9 | 82939.26 | 227.68 | 33.82 | 100.28 | 113.38 | 20.32 | 4.525 | 174.48 | 46.075 | 37.04 | 147.42 | 22.78 | 13.22 | 27.9 | 42.52 | 49.24 | 64.06 | 910.66 | 26.64 | <LOD | <LOD | <LOD | <LOD | 26.975 | <LOD | 731360.22
|  | NIST 679 | 2023-07-18 | stdev | <LOD | 7152.7844 | 14811.4053 | <LOD | <LOD | 2859.4023 | 36.5635 | 743.2635 | 20.7972 | 22.448 | 282.279 | 14597.7759 | 42.0427 | 15.2017 | 17.058 | 23.9329 | 6.3661 | 1.3124 | 37.2129 | 4.3323 | 7.7842 | 34.4644 | 8.65 | 5.2713 | 12.9879 | 24.7256 | 23.3439 | 18.3861 | 342.5361 | 13.3525 | <LOD | <LOD | <LOD | <LOD | 14.2762 | <LOD | 36743.4662
|  | NIST 679 | 2023-07-18 | RSD | <LOD | 8.403% | 7.059% | <LOD | <LOD | 13.04% | 2.439% | 15.36% | 21.48% | 32.36% | 19.14% | 17.6% | 18.47% | 44.95% | 17.01% | 21.11% | 31.33% | 29% | 21.33% | 9.403% | 21.02% | 23.38% | 37.97% | 39.87% | 46.55% | 58.15% | 47.41% | 28.7% | 37.61% | 50.12% | <LOD | <LOD | <LOD | <LOD | 52.92% | <LOD | 5.024%
12 [Col] | 2023-07-18 | SARM 69 | std | rep 1 | 11453.4 | 76044 | 263795.4 | 152.8 | <LOD | 17074.8 | 13967.1 | 4780.2 | 151.8 | 200.4 | 954.4 | 51304.9 | 135.3 | 45.8 | 130.7 | 63.3 | 9.4 | 4.8 | 68.3 | 95.2 | 35.3 | 288.1 | 13.3 | 4.5 | 18.5 | 31.9 | 70.9 | 49 | <LOD | 13.5 | <LOD | 8.1 | <LOD | <LOD | <LOD | 4.6 | 714734.4
13 [Col] | 2023-07-18 | SARM 69 | std | rep 2 | 11444.4 | 77798.8 | 265877.3 | 155.6 | <LOD | 17301.9 | 14126.7 | 4518.7 | 144.5 | 173.8 | 958.8 | 51046 | 344.9 | 39.5 | 123.7 | 65.9 | 14.4 | 3.9 | 71.5 | 95.1 | 30 | 263.7 | 8.6 | 7.4 | 34.4 | 44.3 | 42 | 33.7 | 1241.5 | 31.2 | <LOD | <LOD | <LOD | <LOD | <LOD | 8.9 | 710375.1
14 [Col] | 2023-07-18 | SARM 69 | std | rep 3 | 11980.9 | 74935.1 | 264070.2 | 90.3 | <LOD | 16609.1 | 13955.4 | 4525.7 | 104.1 | 189.2 | 923.5 | 50425.8 | 70.6 | 55.6 | 127 | 56.8 | 9.5 | 4.1 | 71.5 | 98.5 | 31.7 | 290.5 | <LOD | 23.6 | 28 | 37.9 | 47.1 | 43.8 | 1894.9 | <LOD | <LOD | <LOD | <LOD | <LOD | 9.2 | <LOD | 714508.1
15 [Col] | 2023-07-18 | SARM 69 | std | rep 4 | 11664.2 | 77358.8 | 268069 | 157.3 | <LOD | 17333.2 | 14237.9 | 4495.1 | 171.7 | 115.6 | 962.5 | 50623 | 158.3 | 58 | 98.4 | 58.3 | 4.3 | 4.4 | 77.2 | 107.5 | 29.5 | 264 | 10 | 9.7 | <LOD | 36.9 | 51.1 | 47.1 | 456.3 | 57.9 | <LOD | <LOD | <LOD | <LOD | 15.6 | <LOD | 709786.8
16 [Col] | 2023-07-18 | SARM 69 | std | rep 5 | 12951 | 78952.1 | 265823 | 142.6 | <LOD | 17384.4 | 13846.8 | 4592.8 | 124.3 | 197.5 | 845.2 | 51522.2 | 121.5 | 59 | 89.9 | 65.3 | 8.5 | <LOD | 76.8 | 101.7 | 30.5 | 419 | 9.4 | 11 | <LOD | 17.7 | 80.4 | 118 | 1149.7 | 37.1 | <LOD | <LOD | 16.8 | 18.6 | <LOD | <LOD | 707548.8
|  | SARM 69 | 2023-07-18 | average | 11898.78 | 77017.76 | 265526.98 | 139.72 | <LOD | 17140.68 | 14026.78 | 4582.5 | 139.28 | 175.3 | 928.88 | 50984.38 | 166.12 | 51.58 | 113.94 | 61.92 | 9.22 | 4.3 | 73.06 | 99.6 | 31.4 | 305.06 | 10.325 | 11.24 | 26.9667 | 33.74 | 58.3 | 58.32 | 1185.6 | 34.925 | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | 711390.64
|  | SARM 69 | 2023-07-18 | stdev | 627.2485 | 1561.5444 | 1717.032 | 28.2088 | <LOD | 319.9112 | 154.6005 | 116.3259 | 25.9667 | 34.9371 | 49.2648 | 457.8175 | 104.9861 | 8.5406 | 18.4806 | 4.1379 | 3.594 | 0.39158 | 3.8292 | 5.1875 | 2.3281 | 64.9587 | 2.0646 | 7.3378 | 8.0002 | 9.9934 | 16.5086 | 33.88 | 588.6138 | 18.3074 | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | 3132.9956
|  | SARM 69 | 2023-07-18 | RSD | 5.272% | 2.028% | 0.6467% | 20.19% | <LOD | 1.866% | 1.102% | 2.538% | 18.64% | 19.93% | 5.304% | 0.898% | 63.2% | 16.56% | 16.22% | 6.683% | 38.98% | 9.106% | 5.241% | 5.208% | 7.414% | 21.29% | 20% | 65.28% | 29.67% | 29.62% | 28.32% | 58.09% | 49.65% | 52.42% | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | 0.4404%
2 [Col] | 2023-07-28 | NIST 2711a | std | rep 1 | 14363.5 | 69069.9 | 265403.5 | <LOD | <LOD | 26057 | 23132.9 | 3397.4 | 85.4 | 80 | 699.2 | 28811.4 | 66.2 | 26.4 | 172 | 370.6 | 148.8 | 2.9 | 119.2 | 263.8 | 30.8 | 266.4 | 21.5 | 13.1 | 14.3 | 93.1 | 21.6 | 67.3 | 1066.1 | <LOD | <LOD | 19 | 1298.7 | <LOD | 19 | <LOD | 725815
3 [Col] | 2023-07-28 | NIST 2711a | std | rep 2 | 11917.4 | 67136.1 | 259771.4 | <LOD | <LOD | 25097.9 | 23342.1 | 3270.3 | 84.2 | <LOD | 723.1 | 28588.1 | 120.4 | 18.8 | 208 | 397 | 89 | 7.6 | 125.1 | 262.4 | 25.8 | 329.9 | 25.8 | 5.4 | 29.1 | 106.9 | 44.7 | 104.2 | 1057.6 | 34 | <LOD | 8.5 | 1412.9 | <LOD | 11.6 | <LOD | 735882.3
4 [Col] | 2023-07-28 | NIST 2711a | std | rep 3 | 12926.4 | 69693.2 | 269941.9 | <LOD | <LOD | 26351.3 | 23012.4 | 3636.4 | 132.1 | 47.4 | 690.1 | 29321.3 | 138.5 | 19.8 | 205 | 398.9 | 130.2 | 2.4 | 121.7 | 268.1 | 38.2 | 252.8 | 24.2 | 15.6 | 17.9 | 80 | 66.8 | 106.2 | 916.8 | <LOD | <LOD | 10.1 | 1395.7 | <LOD | 10 | 7.4 | 721803
5 [Col] | 2023-07-28 | NIST 2711a | std | rep 4 | 13501.8 | 67772.7 | 260656.6 | <LOD | <LOD | 25619 | 23480.3 | 3790.2 | 210.3 | 41.9 | 655.6 | 28837.4 | 99.1 | 27.7 | 219.2 | 402.4 | 133.6 | <LOD | 125.4 | 258.1 | 28.8 | 307.1 | 21.8 | 7 | 22.2 | 82.6 | 40 | 98.6 | 792.8 | 27.1 | <LOD | 9 | 1343.6 | <LOD | 12.4 | <LOD | 731790.6
6 [Col] | 2023-07-28 | NIST 2711a | std | rep 5 | 14297.8 | 70706.6 | 269225.1 | <LOD | <LOD | 25800.2 | 23144.4 | 3501.6 | <LOD | <LOD | 629.3 | 29407.9 | 106.2 | <LOD | 176.7 | 386.9 | 120.4 | 4.1 | 126.1 | 267 | 32.1 | 285.1 | 15 | 6.1 | 27.9 | 76.7 | 69.2 | 80 | 1307.3 | 44.7 | <LOD | <LOD | 1331.2 | <LOD | 12 | <LOD | 720369.8
|  | NIST 2711a | 2023-07-28 | average | 13401.38 | 68875.7 | 264999.7 | <LOD | <LOD | 25785.08 | 23222.42 | 3519.18 | 128 | 56.4333 | 679.46 | 28993.22 | 106.08 | 23.175 | 196.18 | 391.16 | 124.4 | 4.25 | 123.5 | 263.88 | 31.14 | 288.26 | 21.66 | 9.44 | 22.28 | 87.86 | 48.46 | 91.26 | 1028.12 | 35.2667 | <LOD | 11.65 | 1356.42 | <LOD | 13 | <LOD | 727132.14
|  | NIST 2711a | 2023-07-28 | stdev | 1020.6775 | 1440.5756 | 4707.442 | <LOD | <LOD | 472.9094 | 186.4368 | 202.6531 | 59.2264 | 20.5938 | 37.0562 | 353.9268 | 26.9038 | 4.5243 | 20.6856 | 12.857 | 22.2643 | 2.3445 | 2.9436 | 3.9733 | 4.6063 | 30.96 | 4.1229 | 4.6036 | 6.3437 | 12.2879 | 19.8348 | 16.9218 | 192.3673 | 8.8681 | <LOD | 4.9454 | 47.0768 | <LOD | 3.4756 | <LOD | 6597.9236
|  | NIST 2711a | 2023-07-28 | RSD | 7.616% | 2.092% | 1.776% | <LOD | <LOD | 1.834% | 0.8028% | 5.759% | 46.27% | 36.49% | 5.454% | 1.221% | 25.36% | 19.52% | 10.54% | 3.287% | 17.9% | 55.16% | 2.384% | 1.506% | 14.79% | 10.74% | 19.03% | 48.77% | 28.47% | 13.99% | 40.93% | 18.54% | 18.71% | 25.15% | <LOD | 42.45% | 3.471% | <LOD | 26.74% | <LOD | 0.9074%
7 [Col] | 2023-07-28 | NIST 679 | std | rep 1 | <LOD | 91404.6 | 221429 | <LOD | <LOD | 24130 | 1562.5 | 5685.5 | 105.7 | 80.8 | 1585.7 | 93651.3 | 180.5 | 49.4 | 80.1 | 145 | 26.5 | 7.6 | 191.7 | 55.5 | 39.8 | 168.5 | 16.1 | 10.7 | 19.1 | 27.3 | 65.2 | 31.7 | 716 | <LOD | 4.9 | <LOD | <LOD | <LOD | 22.2 | 4.3 | 702739.1
8 [Col] | 2023-07-28 | NIST 679 | std | rep 2 | <LOD | 86250.2 | 213640.4 | <LOD | <LOD | 23174.9 | 1396.9 | 5418.8 | 79.3 | 43.3 | 1684.3 | 89590.2 | 280.3 | 49.3 | 55 | 125.1 | 28.9 | 2.9 | 187.7 | 51.9 | 39.2 | 154 | 25.9 | <LOD | <LOD | 20.4 | 46.8 | 91.6 | 1052 | <LOD | 6.3 | <LOD | <LOD | <LOD | 20 | <LOD | 718582
9 [Col] | 2023-07-28 | NIST 679 | std | rep 3 | <LOD | 80961.1 | 202047.3 | <LOD | <LOD | 21293 | 1325.3 | 5115.5 | 236.2 | 129.3 | 1350 | 79404.3 | <LOD | 47.2 | 72.1 | 119.5 | 27 | 2.5 | 160.4 | 32.5 | 31.4 | 123.9 | 11.5 | 4.3 | <LOD | 60.4 | 52.2 | 59.5 | 339.5 | <LOD | <LOD | <LOD | <LOD | <LOD | 18.4 | <LOD | 746288
10 [Col] | 2023-07-28 | NIST 679 | std | rep 4 | <LOD | 82882.2 | 207972.3 | <LOD | <LOD | 22277.3 | 1384.9 | 5431.9 | 174.3 | 39.1 | 1608.3 | 86798.3 | <LOD | 69.6 | 44 | 107.1 | 18.9 | 4.7 | 166.9 | 35 | 36.8 | 145 | 21.5 | 9.4 | <LOD | 22.5 | 82.5 | 105 | 897 | 32.9 | <LOD | <LOD | <LOD | <LOD | 14.6 | <LOD | 730172
11 [Col] | 2023-07-28 | NIST 679 | std | rep 5 | <LOD | 92709.2 | 220575.4 | <LOD | <LOD | 23891.6 | 1461.4 | 5831.4 | 210.6 | 99.2 | 1676.2 | 91518.3 | 167.1 | 53.9 | 68.4 | 114.1 | 31.6 | 7.3 | 191.5 | 46.4 | 48.4 | 172.8 | 16.9 | 5.9 | 20.4 | 18.8 | 82.2 | 145.3 | <LOD | 37.6 | <LOD | <LOD | <LOD | <LOD | 14.3 | <LOD | 705104.9
|  | NIST 679 | 2023-07-28 | average | <LOD | 86841.46 | 213132.88 | <LOD | <LOD | 22953.36 | 1426.2 | 5496.62 | 161.22 | 78.34 | 1580.9 | 88192.48 | 209.3 | 53.88 | 63.92 | 122.16 | 26.58 | 5 | 179.64 | 44.26 | 39.12 | 152.84 | 18.38 | 7.575 | <LOD | 29.88 | 65.78 | 86.62 | 751.125 | <LOD | <LOD | <LOD | <LOD | <LOD | 17.9 | <LOD | 720577.2
|  | NIST 679 | 2023-07-28 | stdev | <LOD | 5144.3314 | 8276.4116 | <LOD | <LOD | 1175.554 | 90.2243 | 275.3409 | 67.1283 | 38.0975 | 135.8881 | 5522.4374 | 61.8518 | 9.1212 | 14.3627 | 14.3975 | 4.7368 | 2.3875 | 14.8623 | 10.1653 | 6.1556 | 19.6467 | 5.5002 | 2.9792 | <LOD | 17.3582 | 16.5391 | 43.4674 | 306.8519 | <LOD | <LOD | <LOD | <LOD | <LOD | 3.4278 | <LOD | 18129.2314
|  | NIST 679 | 2023-07-28 | RSD | <LOD | 5.924% | 3.883% | <LOD | <LOD | 5.121% | 6.326% | 5.009% | 41.64% | 48.63% | 8.596% | 6.262% | 29.55% | 16.93% | 22.47% | 11.79% | 17.82% | 47.75% | 8.273% | 22.97% | 15.74% | 12.85% | 29.92% | 39.33% | <LOD | 58.09% | 25.14% | 50.18% | 40.85% | <LOD | <LOD | <LOD | <LOD | <LOD | 19.15% | <LOD | 2.516%
12 [Col] | 2023-07-28 | SARM 69 | std | rep 1 | 13925.2 | 78545.8 | 271282.9 | 129.3 | <LOD | 17273.4 | 14552.1 | 4879.2 | 88 | 217.2 | 951.6 | 50217.9 | 112.7 | 62 | 52.9 | 63.9 | 6.5 | 3.6 | 72 | 93.4 | 29.9 | 268.8 | 6.5 | 9 | <LOD | 42.4 | 40.7 | 81.7 | 853.5 | 27.3 | <LOD | <LOD | 18.2 | <LOD | <LOD | <LOD | 702823.9
13 [Col] | 2023-07-28 | SARM 69 | std | rep 2 | 13011.4 | 77452.3 | 273544.7 | 141.8 | <LOD | 17661.7 | 14765 | 4989.5 | 189.3 | 189.3 | 833.8 | 51730.5 | <LOD | 45.1 | 70.7 | 59.8 | 9.5 | 4.1 | 75.7 | 99.2 | 33.5 | 292.4 | 9.1 | 6.7 | 22.9 | 26.2 | 46.6 | 151.1 | 716.8 | 12.9 | <LOD | <LOD | <LOD | 21.8 | <LOD | <LOD | 700829.5
14 [Col] | 2023-07-28 | SARM 69 | std | rep 3 | 12523.9 | 78836.5 | 268195.7 | 152.2 | <LOD | 17804.8 | 14523.1 | 4495 | 162.5 | 197.9 | 900.2 | 51004.9 | 164.2 | 42.9 | 95.6 | 81.8 | 5.7 | 3.9 | 81 | 108.4 | 30.4 | 250 | 5.5 | 14.8 | 32.5 | 35.8 | 48.4 | 76.9 | 925 | <LOD | <LOD | <LOD | 18.5 | <LOD | 9.5 | <LOD | 706193.1
15 [Col] | 2023-07-28 | SARM 69 | std | rep 4 | 13389.7 | 79767.8 | 272812.3 | 178.1 | <LOD | 17688.9 | 14853.1 | 4889.1 | 176.2 | 199.8 | 1035.5 | 51802.3 | <LOD | 44.1 | 91.4 | 69.4 | 5.6 | 6.2 | 77.7 | 106.6 | 29.4 | 282.4 | 7.5 | 4.8 | <LOD | 18.6 | 48.4 | 104.2 | 1069.6 | 13.5 | <LOD | 4.4 | 17.2 | <LOD | <LOD | <LOD | 698691
16 [Col] | 2023-07-28 | SARM 69 | std | rep 5 | 13046.8 | 79526.7 | 275532.9 | 156.4 | <LOD | 17931.8 | 14781.8 | 5100.6 | 290.1 | 221.9 | 955.9 | 51382.4 | 258 | 64.9 | 90.3 | 68.6 | 13.1 | 4.5 | 79.2 | 103.7 | 30.1 | 305.5 | 9.5 | 5.2 | 30.4 | 75.2 | 82.6 | 110.7 | 514.6 | 17.6 | <LOD | <LOD | <LOD | <LOD | 18 | 7.1 | 696930
|  | SARM 69 | 2023-07-28 | average | 13179.4 | 78825.82 | 272273.7 | 151.56 | <LOD | 17672.12 | 14695.02 | 4870.68 | 181.22 | 205.22 | 935.4 | 51227.6 | 178.3 | 51.8 | 80.18 | 68.7 | 8.08 | 4.46 | 77.12 | 102.26 | 30.66 | 279.82 | 7.62 | 8.1 | 28.6 | 39.64 | 53.34 | 104.92 | 815.9 | 17.825 | <LOD | <LOD | 17.9667 | <LOD | <LOD | <LOD | 701093.5
|  | SARM 69 | 2023-07-28 | stdev | 518.6354 | 914.2814 | 2744.9811 | 18.1588 | <LOD | 247.1869 | 147.8173 | 228.3247 | 72.4323 | 13.7672 | 74.6061 | 647.3486 | 73.6691 | 10.7126 | 18.0288 | 8.2819 | 3.2221 | 1.0262 | 3.4622 | 6.0513 | 1.6288 | 21.4187 | 1.6947 | 4.0915 | 5.0468 | 21.8529 | 16.6595 | 29.5415 | 211.1177 | 6.653 | <LOD | <LOD | 0.68069 | <LOD | <LOD | <LOD | 3611.6699
|  | SARM 69 | 2023-07-28 | RSD | 3.935% | 1.16% | 1.008% | 11.98% | <LOD | 1.399% | 1.006% | 4.688% | 39.97% | 6.709% | 7.976% | 1.264% | 41.32% | 20.68% | 22.49% | 12.06% | 39.88% | 23.01% | 4.489% | 5.918% | 5.312% | 7.654% | 22.24% | 50.51% | 17.65% | 55.13% | 31.23% | 28.16% | 25.88% | 37.32% | <LOD | <LOD | 3.789% | <LOD | <LOD | <LOD | 0.5151%
1 [Col] | 2023-07-31 | NIST 2711a | std | rep 1 | 12707.1 | 69973.2 | 270704 | <LOD | <LOD | 26262.8 | 24432.2 | 3213.7 | 80.8 | 31.7 | 607.4 | 29322.6 | 134.5 | 27.2 | 210.5 | 408.1 | 145.8 | 3.6 | 125 | 273.2 | 30.5 | 276.6 | 22.4 | 8 | 36 | 106.8 | <LOD | <LOD | 979.6 | <LOD | <LOD | 18 | 1381.1 | <LOD | 19.5 | <LOD | 720924.6
2 [Col] | 2023-07-31 | NIST 2711a | std | rep 2 | <LOD | 68773.2 | 265099.5 | <LOD | <LOD | 25923.5 | 23518 | 3431.5 | 149.9 | <LOD | 704.5 | 29575.7 | <LOD | 17.2 | 169.5 | 421.6 | 156.4 | 3.6 | 121.5 | 255.9 | 31.8 | 305.4 | 21.9 | 17.3 | 40.6 | 91.5 | <LOD | 37.8 | <LOD | 21.3 | <LOD | 10.8 | 1371.3 | <LOD | 16.5 | 7.4 | 731682.8
3 [Col] | 2023-07-31 | NIST 2711a | std | rep 3 | 13621.3 | 69320.5 | 266486.9 | 43.1 | <LOD | 25504.7 | 23554.3 | 3314.9 | <LOD | 35.3 | 777.2 | 29070 | 146.7 | 18.9 | 193.1 | 377.6 | 136.7 | 4 | 120.1 | 249.9 | 37.8 | 341.5 | 28.1 | 11 | <LOD | 75.8 | 42.9 | 73.4 | 963.8 | 35.9 | <LOD | 11.4 | 1349.7 | <LOD | <LOD | <LOD | 725256.9
4 [Col] | 2023-07-31 | NIST 2711a | std | rep 4 | 14380.6 | 72043.2 | 269929.3 | 61.3 | <LOD | 26112 | 23385 | 3464.4 | 219.1 | 49.6 | 568.4 | 29471.8 | 52.8 | <LOD | 187.1 | 375.7 | 142.2 | 3 | 122.3 | 261 | 22.2 | 285.1 | 20.9 | 4.5 | 17.1 | 102.7 | 75.3 | 63.7 | 1036.1 | <LOD | <LOD | 4.8 | 1293.7 | <LOD | 10.7 | <LOD | 718363
5 [Col] | 2023-07-31 | NIST 2711a | std | rep 5 | 14976.1 | 71722.3 | 268740.2 | <LOD | <LOD | 26303.7 | 24051.2 | 3229.9 | 164.5 | 114.5 | 576.3 | 28453 | <LOD | 28.1 | 176.1 | 396.1 | 121.6 | 1.7 | 116 | 252.7 | 32.6 | 284.3 | 17.7 | 6.2 | 27.7 | 80.4 | 66.2 | 88 | 620.7 | 19.1 | <LOD | 6.1 | 1353.4 | <LOD | 18.2 | <LOD | 720158.5
|  | NIST 2711a | 2023-07-31 | average | 13921.275 | 70366.48 | 268191.98 | 52.2 | <LOD | 26021.34 | 23788.14 | 3330.88 | 153.575 | 57.775 | 646.76 | 29178.62 | 111.3333 | 22.85 | 187.26 | 395.82 | 140.54 | 3.18 | 120.98 | 258.54 | 30.98 | 298.58 | 22.2 | 9.4 | 30.35 | 91.44 | 61.4667 | 65.725 | 900.05 | 25.4333 | <LOD | 10.22 | 1349.84 | <LOD | 16.225 | <LOD | 723277.16
|  | NIST 2711a | 2023-07-31 | stdev | 981.1288 | 1452.319 | 2350.5213 | 12.8693 | <LOD | 325.0919 | 439.941 | 114.1659 | 56.9286 | 38.5987 | 90.8331 | 447.962 | 51.0571 | 5.5979 | 15.9232 | 19.6994 | 12.8007 | 0.90111 | 3.3071 | 9.1746 | 5.6376 | 26.2554 | 3.7709 | 5.0294 | 10.321 | 13.502 | 16.7106 | 21.1265 | 188.8017 | 9.1309 | <LOD | 5.2108 | 33.9194 | <LOD | 3.8828 | <LOD | 5338.1575
|  | NIST 2711a | 2023-07-31 | RSD | 7.048% | 2.064% | 0.8764% | 24.65% | <LOD | 1.249% | 1.849% | 3.428% | 37.07% | 66.81% | 14.04% | 1.535% | 45.86% | 24.5% | 8.503% | 4.977% | 9.108% | 28.34% | 2.734% | 3.549% | 18.2% | 8.793% | 16.99% | 53.5% | 34.01% | 14.77% | 27.19% | 32.14% | 20.98% | 35.9% | <LOD | 50.99% | 2.513% | <LOD | 23.93% | <LOD | 0.7381%
6 [Col] | 2023-07-31 | NIST 679 | std | rep 1 | <LOD | 89984.2 | 217554.2 | <LOD | <LOD | 23425.3 | 1519.1 | 5332.2 | 140.6 | <LOD | 1674.5 | 93830.3 | 188.1 | 36.8 | 83.7 | 121.7 | 23.5 | 5.4 | 181.8 | 45.3 | 38.8 | 163 | 24.5 | 8.9 | 25.1 | 30.9 | 69.8 | 28.2 | 720.3 | 52.3 | <LOD | <LOD | <LOD | <LOD | <LOD | 4.2 | 708046.1
7 [Col] | 2023-07-31 | NIST 679 | std | rep 2 | <LOD | 88735.2 | 217199.4 | <LOD | <LOD | 24070.1 | 1565.6 | 5352.6 | <LOD | 106.8 | 1667.1 | 92355.9 | 265.2 | 43.6 | 60.8 | 104.6 | 17.2 | 2.1 | 186.6 | 52.2 | 45.8 | 168.9 | 22.7 | 20.9 | 53.2 | 57.3 | 71.9 | 118.5 | 1329.7 | 54.8 | <LOD | <LOD | 18.9 | <LOD | 27.8 | 6.7 | 709568.3
8 [Col] | 2023-07-31 | NIST 679 | std | rep 3 | <LOD | 89633.5 | 219551.7 | <LOD | <LOD | 23684 | 1457.3 | 4556 | 72.5 | 52.3 | 1651.1 | 92242.3 | 180.6 | 27.3 | 30.6 | 113.6 | 19.1 | 2.3 | 190 | 47.9 | 35.1 | 152.9 | 21.8 | 7 | 21.5 | <LOD | 58.7 | 142.5 | 1683.3 | 18.7 | <LOD | <LOD | <LOD | <LOD | 29.3 | <LOD | 707894.7
9 [Col] | 2023-07-31 | NIST 679 | std | rep 4 | <LOD | 86618 | 216105.1 | <LOD | <LOD | 23349.8 | 1519 | 5142.6 | 107.5 | 80.4 | 1513.9 | 87108.9 | 375.3 | 28 | 78.5 | 115.8 | 23.1 | 3.2 | 180.1 | 45.1 | 38.7 | 146.1 | 13.3 | 12.6 | 32.5 | 71.6 | 72.3 | 59.4 | 1479.7 | <LOD | <LOD | <LOD | <LOD | <LOD | 32.7 | <LOD | 718165.8
10 [Col] | 2023-07-31 | NIST 679 | std | rep 5 | 11982.6 | 88737.1 | 218700.1 | <LOD | <LOD | 23745.8 | 1447.8 | 6006.9 | <LOD | 80.3 | 1588.8 | 87990.1 | 260.1 | 53 | 48.2 | 125.1 | 11.4 | 4.4 | 182.8 | 53.3 | 40.1 | 150.6 | 22.3 | 20.6 | 31.5 | 20.1 | 72.8 | 65.5 | <LOD | <LOD | <LOD | <LOD | 15.1 | <LOD | 15 | <LOD | 711190.1
|  | NIST 679 | 2023-07-31 | average | <LOD | 88741.6 | 217822.1 | <LOD | <LOD | 23655 | 1501.76 | 5278.06 | 106.8667 | 79.95 | 1619.08 | 90705.5 | 253.86 | 37.74 | 60.36 | 116.16 | 18.86 | 3.48 | 184.26 | 48.76 | 39.7 | 156.3 | 20.92 | 14 | 32.76 | 44.975 | 69.1 | 82.82 | 1303.25 | 41.9333 | <LOD | <LOD | <LOD | <LOD | 26.2 | <LOD | 710973
|  | NIST 679 | 2023-07-31 | stdev | <LOD | 1308.5563 | 1338.8442 | <LOD | <LOD | 286.0519 | 48.8923 | 519.4745 | 34.0544 | 22.2544 | 67.7906 | 2964.7461 | 78.4325 | 10.8622 | 21.8377 | 7.9204 | 4.949 | 1.4061 | 3.9972 | 3.826 | 3.8839 | 9.3774 | 4.3797 | 6.4834 | 12.2975 | 23.6479 | 5.9249 | 46.5473 | 414.7702 | 20.1594 | <LOD | <LOD | <LOD | <LOD | 7.743 | <LOD | 4236.7252
|  | NIST 679 | 2023-07-31 | RSD | <LOD | 1.475% | 0.6147% | <LOD | <LOD | 1.209% | 3.256% | 9.842% | 31.87% | 27.84% | 4.187% | 3.269% | 30.9% | 28.78% | 36.18% | 6.819% | 26.24% | 40.4% | 2.169% | 7.847% | 9.783% | 6% | 20.94% | 46.31% | 37.54% | 52.58% | 8.574% | 56.2% | 31.83% | 48.07% | <LOD | <LOD | <LOD | <LOD | 29.55% | <LOD | 0.5959%
11 [Col] | 2023-07-31 | SARM 69 | std | rep 1 | <LOD | 79809.1 | 273396.7 | 174.4 | <LOD | 17892.1 | 14871.3 | 4675.6 | 285.1 | 174.8 | 951.5 | 51263 | 195.6 | 57.4 | 94.8 | 80.4 | 6.7 | 3.5 | 78.3 | 99 | 32.5 | 301.8 | 7.1 | 13.2 | 31.5 | 40.1 | 33.4 | 52.2 | 416.9 | <LOD | <LOD | <LOD | <LOD | 25.5 | <LOD | <LOD | 703392
12 [Col] | 2023-07-31 | SARM 69 | std | rep 2 | 12466.3 | 79839.8 | 275590.9 | 214.6 | <LOD | 18162.1 | 15119.5 | 4471.6 | 156.1 | 85.4 | 1021.8 | 52636.5 | 92.3 | 55.5 | 68.1 | 69 | 7 | 4.5 | 73.9 | 99.7 | 29.8 | 307.8 | 6.7 | 15 | 25.5 | 48 | 45.4 | 69.8 | 940.2 | 18.9 | <LOD | <LOD | 15.1 | <LOD | 19.7 | 6 | 696404.9
13 [Col] | 2023-07-31 | SARM 69 | std | rep 3 | 11503.8 | 79102.5 | 273846.4 | 183.9 | <LOD | 17480.9 | 15017 | 4695 | 358.6 | 116.2 | 1012 | 49654 | 309.9 | 57.9 | 64.7 | 68.3 | 7 | 4 | 77.2 | 100.6 | 30.6 | 240.8 | 17 | 12.3 | <LOD | 28.7 | 34.1 | 51.1 | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | 7.8 | <LOD | 703627.4
14 [Col] | 2023-07-31 | SARM 69 | std | rep 4 | 12680.5 | 77637.8 | 268016.8 | 200.7 | <LOD | 16973.2 | 14587 | 4004.3 | 231.3 | 253.7 | 918.7 | 48230.5 | 203.1 | 40 | 59.2 | 57.4 | 3.8 | 3.3 | 69.8 | 112.8 | 28.7 | 257.6 | 9.9 | 14.6 | 20.1 | 36.5 | 46.5 | 92.1 | 502.8 | 17.9 | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | 711133.7
15 [Col] | 2023-07-31 | SARM 69 | std | rep 5 | <LOD | 52187 | 183229.4 | 51.2 | <LOD | 9859.6 | 6761.3 | <LOD | 41.2 | 59.4 | 409.9 | 22205.3 | <LOD | <LOD | 26.9 | 22.7 | 1.2 | 2.6 | 30.6 | <LOD | 10.4 | 66.5 | <LOD | 5.8 | <LOD | <LOD | <LOD | 23.6 | 483.5 | 6.6 | <LOD | <LOD | <LOD | <LOD | 5 | <LOD | 851772.3
|  | SARM 69 | 2023-07-31 | average | 12216.8667 | 73715.24 | 254816.04 | 164.96 | <LOD | 16073.58 | 13271.22 | 4461.625 | 214.46 | 137.9 | 862.78 | 44797.86 | 200.225 | 52.7 | 62.74 | 59.56 | 5.14 | 3.58 | 65.96 | 103.025 | 26.4 | 234.9 | 10.175 | 12.18 | 25.7 | 38.325 | 39.85 | 57.76 | 585.85 | 14.4667 | <LOD | <LOD | <LOD | <LOD | 10.8333 | <LOD | 733266.06
|  | SARM 69 | 2023-07-31 | stdev | 626.7523 | 12067.7207 | 40118.2547 | 65.4421 | <LOD | 3502.5828 | 3644.67 | 321.193 | 121.9373 | 77.7185 | 256.7355 | 12738.0453 | 88.8933 | 8.5296 | 24.2768 | 22.1536 | 2.5822 | 0.71903 | 20.042 | 6.5495 | 9.0512 | 98.3599 | 4.7675 | 3.7272 | 5.7026 | 8.0151 | 7.0638 | 25.3297 | 239.0822 | 6.8311 | <LOD | <LOD | <LOD | <LOD | 7.8053 | <LOD | 66451.5331
|  | SARM 69 | 2023-07-31 | RSD | 5.13% | 16.37% | 15.74% | 39.67% | <LOD | 21.79% | 27.46% | 7.199% | 56.86% | 56.36% | 29.76% | 28.43% | 44.4% | 16.19% | 38.69% | 37.2% | 50.24% | 20.08% | 30.39% | 6.357% | 34.29% | 41.87% | 46.86% | 30.6% | 22.19% | 20.91% | 17.73% | 43.85% | 40.81% | 47.22% | <LOD | <LOD | <LOD | <LOD | 72.05% | <LOD | 9.062%
2 [Col] | 2023-08-01 | NIST 2711a | std | rep 1 | 11843.7 | 68700.7 | 264254.4 | <LOD | <LOD | 25573.2 | 23266.5 | 3299.8 | <LOD | 64.5 | 697.1 | 28108.3 | <LOD | 25.6 | 152 | 401.2 | 143.9 | <LOD | 117.8 | 251.6 | 31.4 | 332.6 | 14.5 | 4.1 | 16.3 | 81.8 | 78.5 | 82.3 | 1652.1 | <LOD | 6.9 | 13 | 1289 | 33.5 | <LOD | <LOD | 730576.2
3 [Col] | 2023-08-01 | NIST 2711a | std | rep 2 | <LOD | 70251.3 | 267427.2 | <LOD | <LOD | 25509.2 | 23212 | <LOD | 98.4 | 111.2 | 597.4 | 28949.6 | <LOD | 19.5 | 177.7 | 413.3 | 137.9 | 4.6 | 124.2 | 260.5 | 39.7 | 303.7 | 19.5 | 10.7 | 32.6 | 109.9 | <LOD | 29.2 | 1582.9 | <LOD | <LOD | <LOD | 1324.2 | <LOD | 10 | 6 | 728713.1
4 [Col] | 2023-08-01 | NIST 2711a | std | rep 3 | 13338.4 | 71830 | 264324.8 | <LOD | <LOD | 25853.1 | 23069.1 | 3686.3 | 163.8 | 122.5 | 669.6 | 28173.8 | 160.5 | 25.1 | 173.6 | 381.6 | 124.8 | <LOD | 114.7 | 257.7 | 30.5 | 279 | 15.3 | 6.8 | 20.8 | 69.2 | <LOD | 53.8 | 977.5 | <LOD | 5.7 | <LOD | 1303.2 | <LOD | <LOD | <LOD | 726343.8
5 [Col] | 2023-08-01 | NIST 2711a | std | rep 4 | 15708.5 | 70646.7 | 267313.9 | <LOD | <LOD | 26339.3 | 23682.4 | 3502.5 | 122.3 | 46 | 728 | 28834 | 89.2 | 22.6 | 179.5 | 413.6 | 136.6 | 5.4 | 117.7 | 261.6 | 44.5 | 322.4 | 23.6 | 7.7 | 33.7 | 91.2 | 44.3 | 47.4 | 1019.5 | <LOD | <LOD | <LOD | 1369.1 | 21.1 | <LOD | <LOD | 720410.7
6 [Col] | 2023-08-01 | NIST 2711a | std | rep 5 | 12120.7 | 71407.3 | 268031.3 | 64.2 | <LOD | 26422 | 24401.2 | 3912 | 150.1 | 58.2 | 679.7 | 29438 | 101.4 | 25.9 | 160.6 | 422.5 | 147.8 | <LOD | 122.8 | 267.4 | 34.8 | 334.3 | 21.2 | 5.5 | <LOD | 43.8 | 47.5 | 46.2 | 390.8 | 42.5 | <LOD | 8.4 | 1353.2 | <LOD | 18.3 | <LOD | 722389.5
|  | NIST 2711a | 2023-08-01 | average | 13252.825 | 70567.2 | 266270.32 | <LOD | <LOD | 25939.36 | 23526.24 | 3600.15 | 133.65 | 80.48 | 674.36 | 28700.74 | 117.0333 | 23.74 | 168.68 | 406.44 | 138.2 | <LOD | 119.44 | 259.76 | 36.18 | 314.4 | 18.82 | 6.96 | 25.85 | 79.18 | 56.7667 | 51.78 | 1124.56 | <LOD | <LOD | <LOD | 1327.74 | <LOD | <LOD | <LOD | 725686.66
|  | NIST 2711a | 2023-08-01 | stdev | 1761.1552 | 1213.5039 | 1828.7607 | <LOD | <LOD | 424.1005 | 539.7058 | 261.035 | 29.1605 | 34.0956 | 48.3952 | 559.4401 | 38.1343 | 2.7043 | 11.8969 | 15.8172 | 8.7587 | <LOD | 3.9412 | 5.7691 | 5.887 | 23.232 | 3.8739 | 2.4936 | 8.6389 | 24.7306 | 18.8895 | 19.3373 | 514.646 | <LOD | <LOD | <LOD | 33.4454 | <LOD | <LOD | <LOD | 4248.6995
|  | NIST 2711a | 2023-08-01 | RSD | 13.29% | 1.72% | 0.6868% | <LOD | <LOD | 1.635% | 2.294% | 7.251% | 21.82% | 42.37% | 7.176% | 1.949% | 32.58% | 11.39% | 7.053% | 3.892% | 6.338% | <LOD | 3.3% | 2.221% | 16.27% | 7.389% | 20.58% | 35.83% | 33.42% | 31.23% | 33.28% | 37.35% | 45.76% | <LOD | <LOD | <LOD | 2.519% | <LOD | <LOD | <LOD | 0.5855%
7 [Col] | 2023-08-01 | NIST 679 | std | rep 1 | <LOD | 93168.4 | 221370.9 | <LOD | <LOD | 24086.2 | 1523.9 | 5328.1 | 145.2 | 62.8 | 1523.7 | 92412.5 | 172.9 | 56.9 | 57.3 | 144.7 | 21.1 | 4.5 | 187.4 | 50.6 | 36 | 164.2 | 14 | 10.4 | 24.9 | <LOD | 80 | 62.1 | 1011.4 | <LOD | <LOD | 8.1 | <LOD | <LOD | 28.9 | <LOD | 702858.2
8 [Col] | 2023-08-01 | NIST 679 | std | rep 2 | <LOD | 91607.8 | 219989.6 | <LOD | <LOD | 23919.2 | 1474.7 | 5399.4 | 90.7 | 84.8 | 1737.3 | 91561.9 | <LOD | 43.3 | 32 | 123.1 | 27.5 | 3.3 | 182 | 42.3 | 35.9 | 154.1 | 14.6 | 15.4 | <LOD | 25.1 | 40.4 | 54.8 | 1154 | 28 | <LOD | <LOD | <LOD | <LOD | 14.6 | <LOD | 706225.1
9 [Col] | 2023-08-01 | NIST 679 | std | rep 3 | <LOD | 94451.2 | 225959 | <LOD | <LOD | 24593.6 | 1587.4 | 5322.5 | 385 | 62.1 | 1707.1 | 91478.7 | 139.6 | 61.1 | 49.1 | 122.4 | 30.1 | 3.4 | 196.5 | 48.2 | 40 | 160.7 | 14.3 | <LOD | 37.7 | 38.6 | 73.4 | 55.7 | 348.3 | 30.3 | <LOD | <LOD | <LOD | <LOD | 30.4 | <LOD | 698463.5
10 [Col] | 2023-08-01 | NIST 679 | std | rep 4 | <LOD | 92202.6 | 222986.3 | <LOD | <LOD | 24206.9 | 1548.3 | 5005.3 | 197.9 | 67.6 | 1615.2 | 92459.4 | 243.7 | 62.9 | 83.4 | 120.5 | 19.7 | 7.1 | 186.5 | 52.5 | 32.8 | 150 | 21.7 | 18.3 | 26.1 | 45.6 | 53.2 | 55.1 | 1731.9 | <LOD | <LOD | 9.6 | <LOD | <LOD | 20.7 | <LOD | 701383.5
11 [Col] | 2023-08-01 | NIST 679 | std | rep 5 | 12465.8 | 88922.6 | 216343.9 | <LOD | <LOD | 23753.1 | 1551.3 | 5339.9 | 130.9 | 75.2 | 1556.9 | 89229.1 | 179.8 | 53.4 | 62.3 | 114.6 | 16.8 | 5.9 | 189.4 | 46.9 | 42.3 | 158.4 | 8.1 | <LOD | <LOD | 16.8 | 66.5 | 105.7 | 1372.8 | 35.2 | <LOD | 7.8 | 16.2 | <LOD | 35.2 | <LOD | 710621.7
|  | NIST 679 | 2023-08-01 | average | <LOD | 92070.52 | 221329.94 | <LOD | <LOD | 24111.8 | 1537.12 | 5279.04 | 189.94 | 70.5 | 1628.04 | 91428.32 | 184 | 55.52 | 56.82 | 125.06 | 23.04 | 4.84 | 188.36 | 48.1 | 37.4 | 157.48 | 14.54 | 14.7 | 29.5667 | 31.525 | 62.7 | 66.68 | 1123.68 | 31.1667 | <LOD | 8.5 | <LOD | <LOD | 25.96 | <LOD | 703910.4
|  | NIST 679 | 2023-08-01 | stdev | <LOD | 2062.3975 | 3564.7759 | <LOD | <LOD | 319.2143 | 41.6088 | 156.0607 | 115.5938 | 9.5504 | 92.6024 | 1312.3154 | 43.4983 | 7.7674 | 18.7909 | 11.4766 | 5.5586 | 1.6426 | 5.2956 | 3.8955 | 3.7463 | 5.5621 | 4.8242 | 3.9962 | 7.0692 | 12.9906 | 15.9355 | 22.0189 | 511.5498 | 3.6774 | <LOD | 0.96437 | <LOD | <LOD | 8.2227 | <LOD | 4678.6871
|  | NIST 679 | 2023-08-01 | RSD | <LOD | 2.24% | 1.611% | <LOD | <LOD | 1.324% | 2.707% | 2.956% | 60.86% | 13.55% | 5.688% | 1.435% | 23.64% | 13.99% | 33.07% | 9.177% | 24.13% | 33.94% | 2.811% | 8.099% | 10.02% | 3.532% | 33.18% | 27.19% | 23.91% | 41.21% | 25.42% | 33.02% | 45.52% | 11.8% | <LOD | 11.35% | <LOD | <LOD | 31.67% | <LOD | 0.6647%
12 [Col] | 2023-08-01 | SARM 69 | std | rep 1 | 11314 | 75952.4 | 262828.5 | 134.6 | <LOD | 17176.5 | 14046.4 | 4530.5 | 103.7 | 161.2 | 949.6 | 51165.8 | 244.2 | 50.1 | 61.3 | 47.1 | 6.2 | 3.5 | 74.1 | 108.6 | 35.8 | 329.7 | 11.8 | 6.7 | 14.5 | 20.9 | 40.2 | 58.8 | 1787.5 | 56.4 | <LOD | <LOD | 16.8 | 35.2 | <LOD | <LOD | 714367.8
13 [Col] | 2023-08-01 | SARM 69 | std | rep 2 | 13411.2 | 79157.7 | 268087.8 | 83.9 | <LOD | 17563.2 | 14462.4 | 5371.1 | 93.1 | 198.5 | 993.5 | 52934.1 | 109.3 | 49 | 73.3 | 83.7 | 5.7 | 1.9 | 78.4 | 114.8 | 35.2 | 313.1 | <LOD | 5.9 | <LOD | 40.2 | 61.1 | 61.8 | 930.1 | 19.8 | <LOD | <LOD | 15.9 | 25.2 | <LOD | <LOD | 702356.9
14 [Col] | 2023-08-01 | SARM 69 | std | rep 3 | 12100.2 | 78886.3 | 271695.6 | 158.2 | <LOD | 17329 | 14509 | 4670.8 | 181.8 | 108.1 | 1033.6 | 51405.9 | 79.9 | 59.6 | 61.2 | 56.4 | 10.5 | 5.6 | 81.3 | 112.6 | 34.4 | 268.2 | 10.1 | 17.4 | 12.6 | 44 | 46.3 | 47.5 | 712.5 | 24.1 | <LOD | <LOD | <LOD | 52.1 | <LOD | 5.4 | 703324.5
15 [Col] | 2023-08-01 | SARM 69 | std | rep 4 | 11440.9 | 77238.6 | 270726.4 | 116.4 | <LOD | 17827.7 | 14660.7 | 4715.9 | 117.1 | 251.7 | 925.6 | 51576.4 | <LOD | 67.3 | 84.5 | 63.1 | 7.5 | 3.6 | 79.2 | 97.4 | 31.1 | 271.4 | 12.9 | 11.6 | <LOD | 28.8 | 52.8 | <LOD | 664.9 | 40.7 | <LOD | <LOD | <LOD | <LOD | 7.8 | <LOD | 705949.6
16 [Col] | 2023-08-01 | SARM 69 | std | rep 5 | 13479.1 | 79346.6 | 271069.4 | 217.7 | <LOD | 17895.5 | 14634.7 | 4181.8 | 194.9 | 164.7 | 885.3 | 51153.6 | 62.6 | 32.4 | 58.7 | 58.9 | 10.6 | 3.3 | 77.9 | 93.3 | 35.1 | 259 | 9.3 | 16.3 | <LOD | 43 | 32.5 | 42.6 | 1277.1 | 32.2 | <LOD | 6.8 | <LOD | <LOD | 9.7 | <LOD | 701909.2
|  | SARM 69 | 2023-08-01 | average | 12349.08 | 78116.32 | 268881.54 | 142.16 | <LOD | 17558.38 | 14462.64 | 4694.02 | 138.12 | 176.84 | 957.52 | 51647.16 | 124 | 51.68 | 67.8 | 61.84 | 8.1 | 3.58 | 78.18 | 105.34 | 34.32 | 288.28 | 11.025 | 11.58 | <LOD | 35.38 | 46.58 | 52.675 | 1074.42 | 34.64 | <LOD | <LOD | <LOD | 37.5 | <LOD | <LOD | 705581.6
|  | SARM 69 | 2023-08-01 | stdev | 1044.4097 | 1470.0232 | 3653.0292 | 50.1959 | <LOD | 310.1024 | 247.1019 | 432.5978 | 46.8649 | 52.8922 | 57.8361 | 740.7322 | 82.4195 | 13.1159 | 10.9289 | 13.555 | 2.3313 | 1.3217 | 2.6243 | 9.4978 | 1.8674 | 31.1332 | 1.6276 | 5.2969 | <LOD | 10.1065 | 11.0516 | 9.1117 | 466.1232 | 14.5559 | <LOD | <LOD | <LOD | 13.5967 | <LOD | <LOD | 5155.3964
|  | SARM 69 | 2023-08-01 | RSD | 8.457% | 1.882% | 1.359% | 35.31% | <LOD | 1.766% | 1.709% | 9.216% | 33.93% | 29.91% | 6.04% | 1.434% | 66.47% | 25.38% | 16.12% | 21.92% | 28.78% | 36.92% | 3.357% | 9.016% | 5.441% | 10.8% | 14.76% | 45.74% | <LOD | 28.57% | 23.73% | 17.3% | 43.38% | 42.02% | <LOD | <LOD | <LOD | 36.26% | <LOD | <LOD | 0.7307%
2 [Col] | 2023-08-02 | NIST 2711a | std | rep 1 | 14904.3 | 72373.6 | 267316.1 | 63.9 | <LOD | 25672.6 | 23292.2 | 3755.4 | 160.3 | <LOD | 721.4 | 29710.7 | 93.7 | 21.5 | 195.3 | 408.8 | 125.7 | 2.4 | 117.5 | 272.3 | 31.4 | 321.6 | 19.3 | 11 | 35.7 | 84.2 | 75.9 | 106.4 | 633.6 | <LOD | <LOD | 13.7 | 1324.3 | <LOD | 10.7 | <LOD | 719786.7
3 [Col] | 2023-08-02 | NIST 2711a | std | rep 2 | 13846 | 68572 | 265496.3 | <LOD | <LOD | 25726.5 | 22985.6 | 3413.4 | 128.7 | <LOD | 562.1 | 28755 | 75.5 | <LOD | 158.1 | 404.2 | 129.8 | 4.7 | 123.2 | 266.5 | 32.8 | 303.3 | 21.8 | 8.7 | 15.6 | 96.9 | 38.3 | 62.9 | 1701.8 | <LOD | 5.9 | <LOD | 1295.3 | 26.1 | <LOD | <LOD | 726537.3
4 [Col] | 2023-08-02 | NIST 2711a | std | rep 3 | 15486.9 | 69826.7 | 265869.5 | <LOD | <LOD | 25236.5 | 23313.5 | 4110.9 | 133.4 | 64.2 | 642.3 | 28957.6 | 127.5 | 16.9 | 167 | 394.5 | 93.4 | 3.3 | 109.7 | 270.9 | 38.6 | 299.7 | 22.4 | <LOD | 16.2 | 90.8 | 28.5 | 88.5 | 838 | <LOD | 6.2 | 7.8 | 1301.8 | <LOD | <LOD | <LOD | 723147.2
5 [Col] | 2023-08-02 | NIST 2711a | std | rep 4 | 13166 | 70375.6 | 273831.8 | <LOD | <LOD | 26689.4 | 23746.1 | 3861.6 | <LOD | <LOD | 657.4 | 30413.2 | 103.3 | 22.7 | 177.2 | 397.3 | 169.2 | 3.6 | 117.3 | 261.7 | 30.3 | 288.8 | 18.4 | 15.8 | <LOD | 93.2 | <LOD | 80.6 | 2215.8 | 44.7 | <LOD | 15.2 | 1365.5 | <LOD | 20.5 | 8.1 | 714400.6
6 [Col] | 2023-08-02 | NIST 2711a | std | rep 5 | 13497.5 | 70781.5 | 270252.5 | <LOD | <LOD | 26623.9 | 23924.9 | 3630.2 | 204.6 | 34.7 | 606.2 | 29020.2 | <LOD | 15.1 | 154.6 | 397.8 | 127.9 | 6.4 | 116.5 | 261.8 | 34.4 | 284.7 | 23 | 6.6 | 27.8 | 104.7 | 38.8 | 56 | 1025.7 | 44.4 | <LOD | <LOD | 1320 | <LOD | <LOD | <LOD | 719794.5
|  | NIST 2711a | 2023-08-02 | average | 14180.14 | 70385.88 | 268553.24 | <LOD | <LOD | 25989.78 | 23452.46 | 3754.3 | 156.75 | <LOD | 637.88 | 29371.34 | 100 | 19.05 | 170.44 | 400.52 | 129.2 | 4.08 | 116.84 | 266.64 | 33.5 | 299.62 | 20.98 | 10.525 | 23.825 | 93.96 | 45.375 | 78.88 | 1282.98 | <LOD | <LOD | 12.2333 | 1321.38 | <LOD | <LOD | <LOD | 720733.26
|  | NIST 2711a | 2023-08-02 | stdev | 979.5606 | 1388.3889 | 3493.871 | <LOD | <LOD | 638.147 | 378.2775 | 259.9479 | 34.8054 | <LOD | 59.4188 | 684.6518 | 21.657 | 3.6309 | 16.422 | 5.8333 | 26.907 | 1.5353 | 4.8008 | 4.9506 | 3.2388 | 14.4598 | 2.0154 | 3.9492 | 9.7058 | 7.5798 | 20.8952 | 20.1972 | 657.9244 | <LOD | <LOD | 3.9119 | 27.4798 | <LOD | <LOD | <LOD | 4510.9261
|  | NIST 2711a | 2023-08-02 | RSD | 6.908% | 1.973% | 1.301% | <LOD | <LOD | 2.455% | 1.613% | 6.924% | 22.2% | <LOD | 9.315% | 2.331% | 21.66% | 19.06% | 9.635% | 1.456% | 20.83% | 37.63% | 4.109% | 1.857% | 9.668% | 4.826% | 9.606% | 37.52% | 40.74% | 8.067% | 46.05% | 25.6% | 51.28% | <LOD | <LOD | 31.98% | 2.08% | <LOD | <LOD | <LOD | 0.6259%
7 [Col] | 2023-08-02 | NIST 679 | std | rep 1 | <LOD | 91872.5 | 223046.5 | <LOD | <LOD | 24141 | 1455 | 5173.9 | 197.3 | 158.7 | 1442.5 | 89659.9 | 373.7 | 60.8 | 74.5 | 124.2 | 30.5 | 2.9 | 189.4 | 52.1 | 42.5 | 160.1 | 8.4 | 7 | <LOD | 34.7 | 97.5 | 130.9 | 1026.5 | 26.7 | <LOD | <LOD | <LOD | <LOD | 27.4 | <LOD | 704868.4
8 [Col] | 2023-08-02 | NIST 679 | std | rep 2 | <LOD | 90330.7 | 218594.9 | <LOD | <LOD | 24268.3 | 1559.9 | 5656.4 | 84.7 | 40.8 | 1627.5 | 93097.5 | 96 | 78.7 | 39.2 | 119.3 | 14 | 5.7 | 189.9 | 45.1 | 33.8 | 157.7 | 23.6 | 10.4 | 23.2 | 53.4 | 55.4 | 108.5 | 913.4 | 20.4 | <LOD | <LOD | <LOD | 20.5 | <LOD | <LOD | 706575.4
9 [Col] | 2023-08-02 | NIST 679 | std | rep 3 | <LOD | 94513.5 | 227128 | <LOD | <LOD | 24955.6 | 1642.4 | 5744.4 | 117.8 | 98.4 | 1714.4 | 93163.2 | 211 | 50.2 | 104.9 | 125.7 | 18.8 | 2.7 | 197.6 | 49.4 | 36.2 | 167.9 | 27.5 | 13.7 | 26 | 31.7 | 54.2 | 56.5 | 492.6 | 52.1 | <LOD | <LOD | 16.4 | <LOD | 13.5 | <LOD | 694911.3
10 [Col] | 2023-08-02 | NIST 679 | std | rep 4 | <LOD | 91120 | 221946.5 | <LOD | <LOD | 24655 | 1482.5 | 5645.7 | 86.9 | 70.4 | 1648.4 | 91807.2 | <LOD | 23.3 | 80.2 | 137 | 15.6 | 5.5 | 184.5 | 60.7 | 27.2 | 162.8 | 23.8 | 14.2 | 25.5 | 43.6 | 125.4 | 111.4 | <LOD | <LOD | <LOD | 5.5 | 23.5 | <LOD | 35.7 | <LOD | 704864.4
11 [Col] | 2023-08-02 | NIST 679 | std | rep 5 | <LOD | 96929.8 | 227238.5 | <LOD | <LOD | 24690.7 | 1465.6 | 5189.7 | 82 | 52.5 | 1725.4 | 91240.5 | 284.7 | 45.8 | 48.6 | 127.8 | 23.2 | 7.9 | 191.4 | 53.7 | 48.2 | 157.7 | 21.5 | 11.8 | <LOD | 37.3 | 54.5 | 127.1 | 1163.9 | 20.2 | <LOD | <LOD | <LOD | <LOD | 24 | <LOD | 695069.7
|  | NIST 679 | 2023-08-02 | average | <LOD | 92953.3 | 223590.88 | <LOD | <LOD | 24542.12 | 1521.08 | 5482.02 | 113.74 | 84.16 | 1631.64 | 91793.66 | 241.35 | 51.76 | 69.48 | 126.8 | 20.42 | 4.94 | 190.56 | 52.2 | 37.58 | 161.24 | 20.96 | 11.42 | 24.9 | 40.14 | 77.4 | 106.88 | 899.1 | 29.85 | <LOD | <LOD | <LOD | <LOD | 25.15 | <LOD | 701257.84
|  | NIST 679 | 2023-08-02 | stdev | <LOD | 2722.7709 | 3666.6231 | <LOD | <LOD | 332.2869 | 79.3072 | 276.7814 | 48.9129 | 46.9861 | 113.7015 | 1452.1312 | 117.5351 | 20.3434 | 26.2085 | 6.505 | 6.6409 | 2.1698 | 4.7109 | 5.7611 | 8.0766 | 4.2764 | 7.346 | 2.9003 | 1.4933 | 8.6124 | 32.6139 | 29.7784 | 289.7105 | 15.1373 | <LOD | <LOD | <LOD | <LOD | 9.1908 | <LOD | 5763.929
|  | NIST 679 | 2023-08-02 | RSD | <LOD | 2.929% | 1.64% | <LOD | <LOD | 1.354% | 5.214% | 5.049% | 43% | 55.83% | 6.969% | 1.582% | 48.7% | 39.3% | 37.72% | 5.13% | 32.52% | 43.92% | 2.472% | 11.04% | 21.49% | 2.652% | 35.05% | 25.4% | 5.997% | 21.46% | 42.14% | 27.86% | 32.22% | 50.71% | <LOD | <LOD | <LOD | <LOD | 36.54% | <LOD | 0.8219%
12 [Col] | 2023-08-02 | SARM 69 | std | rep 1 | 12911.4 | 80094.2 | 275286.3 | 162.3 | <LOD | 17828.4 | 15269.3 | 4763.9 | 210.4 | 209.2 | 969.9 | 51632.9 | 106.3 | 59.4 | 79.2 | 67.8 | 4.5 | 4.4 | 74.5 | 96 | 35.5 | 283.4 | 5.8 | 10.3 | <LOD | <LOD | <LOD | 66 | 383.7 | 43.1 | <LOD | <LOD | 19.8 | <LOD | <LOD | <LOD | 697367.6
13 [Col] | 2023-08-02 | SARM 69 | std | rep 2 | 11314 | 75451.1 | 267164.3 | 156.5 | <LOD | 17252.8 | 14497.8 | 4491.1 | 277.5 | 119.9 | 993.1 | 49725.5 | 130.5 | 50.6 | 80.2 | 74.3 | 10.3 | 4 | 70.2 | 106 | 31.2 | 306.8 | 11.7 | 12.3 | 12.7 | 23.9 | 53.2 | 25.6 | 790.4 | <LOD | <LOD | 4.9 | <LOD | 28.7 | <LOD | <LOD | 712889.9
14 [Col] | 2023-08-02 | SARM 69 | std | rep 3 | 12876.9 | 79211.8 | 273694.6 | 187.8 | <LOD | 17662.2 | 15221.4 | 5024.1 | 193 | 168.9 | 891.9 | 51612.6 | 314.7 | 81.4 | 64.8 | 66.1 | 10.9 | 3.5 | 78.3 | 104.9 | 31.9 | 291.9 | 11.2 | <LOD | 42.6 | 75.6 | 68.5 | 49.3 | 757.9 | <LOD | <LOD | <LOD | <LOD | 35.8 | <LOD | 3.8 | 698783.6
15 [Col] | 2023-08-02 | SARM 69 | std | rep 4 | 12073.2 | 80963.3 | 275159.3 | 230.7 | <LOD | 17684 | 15415.2 | 4741.6 | 115.1 | 187.1 | 958.7 | 51234.6 | <LOD | 71 | 83.8 | 67.9 | 9 | 2.1 | 73.8 | 104.9 | 33.2 | 267.3 | 6.8 | 12.2 | 26 | 39 | 49.8 | 76.8 | 931.6 | 13.5 | <LOD | <LOD | 15.8 | <LOD | <LOD | <LOD | 697715.3
16 [Col] | 2023-08-02 | SARM 69 | std | rep 5 | 12482.3 | 79668.9 | 272626.5 | 176.6 | <LOD | 17794 | 14868.7 | 4573.6 | 105.4 | 142.6 | 994.2 | 50938 | 342.4 | 58.5 | 77.9 | 64.5 | 4.8 | 3.3 | 74.4 | 98.6 | 29.2 | 262 | 9.1 | 9.6 | 12.5 | <LOD | 54.2 | 66.6 | 1154.2 | 25.6 | <LOD | <LOD | 15.7 | 20.1 | <LOD | <LOD | 700917.4
|  | SARM 69 | 2023-08-02 | average | 12331.56 | 79077.86 | 272786.2 | 182.78 | <LOD | 17644.28 | 15054.48 | 4718.86 | 180.28 | 165.54 | 961.56 | 51028.72 | 223.475 | 64.18 | 77.18 | 68.12 | 7.9 | 3.46 | 74.24 | 102.08 | 32.2 | 282.28 | 8.92 | 11.1 | 23.45 | 46.1667 | 56.425 | 56.86 | 803.56 | 27.4 | <LOD | <LOD | 17.1 | 28.2 | <LOD | <LOD | 701534.76
|  | SARM 69 | 2023-08-02 | stdev | 663.0443 | 2127.7034 | 3329.3602 | 29.4525 | <LOD | 229.9302 | 370.316 | 205.3608 | 71.3726 | 35.3255 | 41.8043 | 783.4137 | 122.2559 | 12.0682 | 7.2596 | 3.7259 | 3.0471 | 0.8735 | 2.8763 | 4.4819 | 2.3441 | 18.2394 | 2.6071 | 1.3589 | 14.2442 | 26.5846 | 8.2674 | 20.0646 | 281.9056 | 14.8819 | <LOD | <LOD | 2.3388 | 7.8619 | <LOD | <LOD | 6496.9746
|  | SARM 69 | 2023-08-02 | RSD | 5.377% | 2.691% | 1.221% | 16.11% | <LOD | 1.303% | 2.46% | 4.352% | 39.59% | 21.34% | 4.348% | 1.535% | 54.71% | 18.8% | 9.406% | 5.47% | 38.57% | 25.25% | 3.874% | 4.391% | 7.28% | 6.461% | 29.23% | 12.24% | 60.74% | 57.58% | 14.65% | 35.29% | 35.08% | 54.31% | <LOD | <LOD | 13.68% | 27.88% | <LOD | <LOD | 0.9261%
2 [Col] | 2023-08-03 | NIST 2711a | std | rep 1 | 13256 | 72252.2 | 271126.3 | <LOD | <LOD | 26104.5 | 24448.5 | <LOD | 197.2 | <LOD | 787.5 | 29307 | 208.9 | 28.5 | 170.9 | 389.4 | 136 | <LOD | 119.4 | 258.7 | 34.7 | 251.9 | 17.4 | 14 | <LOD | 96.9 | 43.4 | 103.7 | 1212.6 | 40.5 | <LOD | <LOD | 1350.2 | <LOD | <LOD | <LOD | 717910.6
3 [Col] | 2023-08-03 | NIST 2711a | std | rep 2 | 13390.4 | 72898 | 272749.4 | <LOD | <LOD | 26431.3 | 24295.8 | 3071.2 | 132.7 | 44.2 | 696.1 | 28402 | 56.3 | 42.9 | 217.2 | 370.1 | 135.5 | <LOD | 118.5 | 263.7 | 41.4 | 353.5 | 13.8 | <LOD | <LOD | 66.3 | 47.9 | 32.8 | 426.6 | <LOD | <LOD | 15.5 | 1310.2 | 39.1 | <LOD | 4.8 | 717526.2
4 [Col] | 2023-08-03 | NIST 2711a | std | rep 3 | 15005.5 | 71096.6 | 271226.3 | <LOD | <LOD | 26485.3 | 24377.8 | 3403.7 | 71.2 | 35.9 | 673.6 | 29044.8 | 46.7 | 27.2 | 208.8 | 381.1 | 138.2 | 1.8 | 115.5 | 275.4 | 33.5 | 341.4 | 16.2 | 14.1 | <LOD | 40.6 | 19.4 | 70.9 | 779.6 | 17.3 | 5.4 | 7.9 | 1385.8 | <LOD | 8.7 | <LOD | 717084.7
5 [Col] | 2023-08-03 | NIST 2711a | std | rep 4 | 13690.3 | 71131.2 | 269403.4 | <LOD | <LOD | 26384.8 | 24049.9 | 3229.5 | 119.7 | 70.8 | 735.7 | 30386.7 | 82.3 | 31.9 | 177.6 | 431.2 | 156 | 3 | 116.5 | 270.1 | 38.6 | 289.2 | 24.9 | 20.5 | 16.2 | 58.2 | 27.6 | 115.4 | 900.7 | <LOD | <LOD | 10.6 | 1331.2 | <LOD | <LOD | <LOD | 719050.8
6 [Col] | 2023-08-03 | NIST 2711a | std | rep 5 | 13943.6 | 71189.6 | 271424.7 | <LOD | <LOD | 26253.2 | 23737.1 | 3500.7 | 92.1 | 45.8 | 588.2 | 28682.2 | 128.6 | <LOD | 180.9 | 396.1 | 136.4 | 7 | 123.4 | 261 | 33.5 | 268.3 | 16.5 | 9 | 14 | 72.4 | 18.1 | 40.4 | 720.9 | <LOD | <LOD | 7 | 1331.8 | 17.1 | 16.5 | <LOD | 719040.5
|  | NIST 2711a | 2023-08-03 | average | 13857.16 | 71713.52 | 271186.02 | <LOD | <LOD | 26331.82 | 24181.82 | 3301.275 | 122.58 | 49.175 | 696.22 | 29164.54 | 104.56 | 32.625 | 191.08 | 393.58 | 140.42 | 3.9333 | 118.66 | 265.78 | 36.34 | 300.86 | 17.76 | 14.4 | <LOD | 66.88 | 31.28 | 72.64 | 808.08 | <LOD | <LOD | 10.25 | 1341.84 | <LOD | <LOD | <LOD | 718122.56
|  | NIST 2711a | 2023-08-03 | stdev | 695.2081 | 819.6566 | 1192.0621 | <LOD | <LOD | 153.3906 | 290.5583 | 190.0415 | 48.0827 | 15.0555 | 74.2811 | 765.2494 | 66.4023 | 7.1309 | 20.5477 | 23.1633 | 8.7688 | 2.7227 | 3.0697 | 6.8649 | 3.5189 | 44.7427 | 4.2075 | 4.7124 | <LOD | 20.6033 | 13.7068 | 36.8199 | 285.5492 | <LOD | <LOD | 3.8197 | 28.3614 | <LOD | <LOD | <LOD | 891.9037
|  | NIST 2711a | 2023-08-03 | RSD | 5.017% | 1.143% | 0.4396% | <LOD | <LOD | 0.5825% | 1.202% | 5.757% | 39.23% | 30.62% | 10.67% | 2.624% | 63.51% | 21.86% | 10.75% | 5.885% | 6.245% | 69.22% | 2.587% | 2.583% | 9.683% | 14.87% | 23.69% | 32.72% | <LOD | 30.81% | 43.82% | 50.69% | 35.34% | <LOD | <LOD | 37.27% | 2.114% | <LOD | <LOD | <LOD | 0.1242%
7 [Col] | 2023-08-03 | NIST 679 | std | rep 1 | <LOD | 94568 | 225547 | <LOD | <LOD | 24619 | 1540.2 | 5839.9 | 320 | 76.3 | 1388.4 | 91467.4 | 217.4 | 38.1 | 62.6 | 125.1 | 15.4 | 5.1 | 190.9 | 62.3 | 48.4 | 158.1 | 10.7 | 4.2 | 27.6 | <LOD | 33.8 | 75.3 | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | 31.6 | 6.9 | 698966.9
8 [Col] | 2023-08-03 | NIST 679 | std | rep 2 | <LOD | 91842 | 220225.3 | <LOD | <LOD | 24001.8 | 1626.8 | 5565.9 | 201 | 126.4 | 1704.1 | 92040.4 | <LOD | 51.9 | 80.7 | 127.4 | 16.9 | 1.8 | 187.7 | 49 | 42.2 | 174.1 | 22.1 | 10.4 | 41.8 | 62.2 | 37.5 | 69.6 | 659.5 | 27 | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | 705234.5
9 [Col] | 2023-08-03 | NIST 679 | std | rep 3 | <LOD | 94413.3 | 226600.7 | <LOD | <LOD | 24768 | 1515.6 | 5338.8 | 244.9 | 71.5 | 1692.8 | 95273.2 | <LOD | 68.8 | 49.2 | 125.7 | 30.1 | 2.6 | 179.8 | 48.1 | 32.9 | 150.7 | 19.6 | 10 | 18.7 | 51.6 | 62.7 | 110.2 | 794.2 | <LOD | <LOD | <LOD | <LOD | <LOD | 53.3 | <LOD | 693840.3
10 [Col] | 2023-08-03 | NIST 679 | std | rep 4 | <LOD | 96429.1 | 228069 | <LOD | <LOD | 25379 | 1679.9 | 5958.9 | 187.6 | 101.6 | 1857.6 | 93325 | <LOD | 35.2 | 87.8 | 123.5 | 24.9 | 3.6 | 201.9 | 48.8 | 38.3 | 165.7 | 18.6 | 13.1 | 37.9 | 18 | 55.1 | 92.3 | <LOD | 55.7 | <LOD | <LOD | <LOD | <LOD | 38.1 | <LOD | 692341.8
11 [Col] | 2023-08-03 | NIST 679 | std | rep 5 | <LOD | 92394.1 | 224719.3 | <LOD | <LOD | 24825.5 | 1563.3 | 5762.1 | 162.7 | 56.6 | 1507.3 | 93611.7 | 174.2 | 39.1 | 60 | 125.5 | 21.4 | 7 | 200.6 | 51.5 | 39.9 | 152.8 | 14 | 10.1 | <LOD | 22.4 | 43.1 | 65.7 | <LOD | 44.3 | <LOD | <LOD | 15.6 | <LOD | 40.1 | 5 | 699270.6
|  | NIST 679 | 2023-08-03 | average | <LOD | 93929.3 | 225032.26 | <LOD | <LOD | 24718.66 | 1585.16 | 5693.12 | 223.24 | 86.48 | 1630.04 | 93143.54 | <LOD | 46.62 | 68.06 | 125.44 | 21.74 | 4.02 | 192.18 | 51.94 | 40.34 | 160.28 | 17 | 9.56 | 31.5 | 38.55 | 46.44 | 82.62 | <LOD | 42.3333 | <LOD | <LOD | <LOD | <LOD | 40.775 | <LOD | 697930.82
|  | NIST 679 | 2023-08-03 | stdev | <LOD | 1844.2379 | 2964.4375 | <LOD | <LOD | 493.3859 | 67.1703 | 244.2689 | 61.7689 | 27.582 | 183.4818 | 1484.412 | <LOD | 13.9588 | 15.8107 | 1.3957 | 5.9919 | 2.0717 | 9.2243 | 5.9324 | 5.6598 | 9.6546 | 4.5831 | 3.2578 | 10.4259 | 21.7006 | 12.1473 | 18.4645 | <LOD | 14.4507 | <LOD | <LOD | <LOD | <LOD | 9.1043 | <LOD | 5103.4722
|  | NIST 679 | 2023-08-03 | RSD | <LOD | 1.963% | 1.317% | <LOD | <LOD | 1.996% | 4.237% | 4.291% | 27.67% | 31.89% | 11.26% | 1.594% | <LOD | 29.94% | 23.23% | 1.113% | 27.56% | 51.54% | 4.8% | 11.42% | 14.03% | 6.024% | 26.96% | 34.08% | 33.1% | 56.29% | 26.16% | 22.35% | <LOD | 34.14% | <LOD | <LOD | <LOD | <LOD | 22.33% | <LOD | 0.7312%
12 [Col] | 2023-08-03 | SARM 69 | std | rep 1 | 13206.4 | 79680.3 | 273349.9 | 160.9 | <LOD | 17850.9 | 14710.9 | 4464.2 | 315.4 | 258.6 | 965.6 | 51543.6 | 290.6 | 59.6 | 69.4 | 72.6 | 6 | 3.3 | 76.7 | 105.6 | 30.8 | 292.9 | 13.8 | 15 | 13.6 | 39.5 | 67.4 | 82.7 | 1199.9 | 16.4 | <LOD | <LOD | 17 | 33 | <LOD | 7 | 698561.9
13 [Col] | 2023-08-03 | SARM 69 | std | rep 2 | 12033.3 | 80845.2 | 277054.4 | 193 | <LOD | 18043.1 | 15100.6 | 5079.4 | 138.1 | 289.9 | 991.6 | 52495.4 | 88.4 | 66.9 | 95.2 | 88.6 | 14 | 3.9 | 79.7 | 96.7 | 33.1 | 341.3 | 6.6 | 5.1 | 24.1 | 37.9 | 57.5 | 131.6 | 439.4 | <LOD | <LOD | 8.3 | <LOD | <LOD | 23.9 | <LOD | 694565.7
14 [Col] | 2023-08-03 | SARM 69 | std | rep 3 | 12111.3 | 79589.8 | 266934 | 168.5 | <LOD | 17500 | 14685.7 | 5180.9 | 101.5 | 271.2 | 971.3 | 52425.6 | 137.4 | 63 | 81 | 77 | 9.3 | 5.9 | 78 | 96.6 | 25.4 | 297.1 | 14.5 | 9.2 | 17.7 | 52.6 | 58.7 | 64.4 | <LOD | 16.6 | <LOD | 6.3 | <LOD | <LOD | 10.1 | <LOD | 705958.6
15 [Col] | 2023-08-03 | SARM 69 | std | rep 4 | 13215.8 | 79105 | 270687 | 173.7 | <LOD | 17935.4 | 15101.3 | 4497.5 | <LOD | 142.7 | 818.6 | 49860.8 | 82.7 | 58.5 | 64.1 | 74.4 | <LOD | 4.3 | 71 | 90 | 26.7 | 306 | 8.6 | 7.8 | 14.3 | 15.4 | 43.2 | 104.8 | 1264.1 | 14.1 | <LOD | <LOD | 19.4 | <LOD | <LOD | 3.3 | 703623.8
16 [Col] | 2023-08-03 | SARM 69 | std | rep 5 | 12103.6 | 80910.9 | 277523.9 | 192.2 | <LOD | 17877.5 | 14888.3 | 4590.6 | 123.8 | 273.1 | 884.8 | 50384.8 | <LOD | 58.4 | 69.5 | 63.9 | 7.1 | <LOD | 74.4 | 104.8 | 28 | 274.2 | 7.9 | 8.4 | <LOD | 26.9 | 39.6 | 37.4 | 423.9 | 23.9 | <LOD | <LOD | 17.9 | 22.8 | <LOD | <LOD | 697389.2
|  | SARM 69 | 2023-08-03 | average | 12534.08 | 80026.24 | 273109.84 | 177.66 | <LOD | 17841.38 | 14897.36 | 4762.52 | 169.7 | 247.1 | 926.38 | 51342.04 | 149.775 | 61.28 | 75.84 | 75.3 | 9.1 | 4.35 | 75.96 | 98.74 | 28.8 | 302.3 | 10.28 | 9.1 | 17.425 | 34.46 | 53.28 | 84.18 | 831.825 | 17.75 | <LOD | <LOD | 18.1 | <LOD | <LOD | <LOD | 700019.84
|  | SARM 69 | 2023-08-03 | stdev | 618.7877 | 808.1083 | 4447.2537 | 14.3806 | <LOD | 204.6132 | 201.5858 | 340.6764 | 98.2939 | 59.4144 | 72.7239 | 1189.0488 | 97.0408 | 3.6534 | 12.4593 | 8.914 | 3.5431 | 1.1121 | 3.3813 | 6.4983 | 3.1265 | 24.6967 | 3.6134 | 3.6401 | 4.7968 | 14.0244 | 11.5684 | 36.2479 | 462.8683 | 4.254 | <LOD | <LOD | 1.2124 | <LOD | <LOD | <LOD | 4665.0535
|  | SARM 69 | 2023-08-03 | RSD | 4.937% | 1.01% | 1.628% | 8.094% | <LOD | 1.147% | 1.353% | 7.153% | 57.92% | 24.04% | 7.85% | 2.316% | 64.79% | 5.962% | 16.43% | 11.84% | 38.93% | 25.56% | 4.451% | 6.581% | 10.86% | 8.17% | 35.15% | 40% | 27.53% | 40.7% | 21.71% | 43.06% | 55.64% | 23.97% | <LOD | <LOD | 6.699% | <LOD | <LOD | <LOD | 0.6664%
2 [Col] | 2023-08-22 | NIST2711a | standard | rep 1 | 14759.7 | 70232 | 268660.8 | <LOD | <LOD | 26393 | 23816.1 | 3677.4 | 82.9 | <LOD | 683.3 | 29995.6 | 132.2 | <LOD | 168.3 | 432.4 | 127.5 | <LOD | 121.1 | 277.9 | 36 | 272.6 | 25.5 | 18.8 | 21.2 | 60.5 | 63.2 | 106.8 | 1269.2 | 17.7 | <LOD | 13.3 | 1394.8 | <LOD | 25 | <LOD | 718910
3 [Col] | 2023-08-22 | NIST2711a | standard | rep 2 | 14426.3 | 69389.7 | 265666.6 | <LOD | <LOD | 25814.6 | 23732.2 | 3057.1 | <LOD | 59.7 | 630.3 | 27818.7 | 165.1 | 19.8 | 167.6 | 402.4 | 115.1 | 5.2 | 120 | 250.9 | 35.6 | 302.4 | 21 | 7.1 | 17.9 | 52.7 | 30.3 | 68.1 | 523.3 | <LOD | <LOD | 13.2 | 1282.4 | <LOD | 12 | <LOD | 727036.1
4 [Col] | 2023-08-22 | NIST2711a | standard | rep 3 | 12754.6 | 71234.9 | 270875.5 | <LOD | <LOD | 26686.1 | 24899.1 | 3341.5 | <LOD | <LOD | 643.2 | 31171 | 107.5 | <LOD | 186.3 | 414.7 | 164.4 | <LOD | 125 | 286.2 | 35.9 | 258.1 | 17 | 7.3 | 15 | 91 | 64.2 | 87.5 | 1554.2 | 39.5 | <LOD | 13.3 | 1383.2 | <LOD | 26.7 | <LOD | 716585.5
5 [Col] | 2023-08-22 | NIST2711a | standard | rep 4 | 13523.7 | 72131.1 | 273188.4 | <LOD | <LOD | 26731.6 | 25128.7 | 3236.9 | 174.4 | 37.2 | 563.6 | 30852 | <LOD | <LOD | 165.8 | 422.9 | 129.4 | 2 | 121 | 263.7 | 27.6 | 322.4 | 24.7 | 13 | 42.3 | 81.5 | 76.7 | 96.7 | 412.5 | 17.4 | <LOD | 15.6 | 1393.1 | <LOD | 7.5 | <LOD | 714227.1
6 [Col] | 2023-08-22 | NIST2711a | standard | rep 5 | 15314.6 | 69374.1 | 268264.5 | <LOD | <LOD | 26715.8 | 23750.7 | 3051.1 | 105.1 | 47.5 | 668 | 28665.7 | <LOD | 28.7 | 178 | 376.6 | 145.4 | 3.2 | 117.2 | 270.2 | 34 | 292.9 | 22.8 | 8.9 | 59.7 | 74.9 | 33.5 | 51.8 | 904.5 | 53.5 | <LOD | <LOD | 1360.8 | <LOD | 8.6 | 8 | 721682
|  | NIST2711a | 2023-08-22 | average | 14155.78 | 70472.36 | 269331.16 | <LOD | <LOD | 26468.22 | 24265.36 | 3272.8 | 120.8 | 48.1333 | 637.68 | 29700.6 | 134.9333 | <LOD | 173.2 | 409.8 | 136.36 | 3.4667 | 120.86 | 269.78 | 33.82 | 289.68 | 22.2 | 11.02 | 31.22 | 72.12 | 53.58 | 82.18 | 932.74 | 32.025 | <LOD | 13.85 | 1362.86 | <LOD | 15.96 | <LOD | 719688.14
|  | NIST2711a | 2023-08-22 | stdev | 1017.7911 | 1200.9357 | 2840.7496 | <LOD | <LOD | 390.8368 | 688.8308 | 257.5786 | 47.7277 | 11.2634 | 46.296 | 1430.6807 | 28.8971 | <LOD | 8.729 | 21.5834 | 19.019 | 1.6166 | 2.7996 | 13.5014 | 3.5696 | 25.1584 | 3.3904 | 4.9535 | 19.2015 | 15.5268 | 20.5245 | 22.1724 | 484.3826 | 17.6649 | <LOD | 1.1676 | 46.9744 | <LOD | 9.1991 | <LOD | 4950.3239
|  | NIST2711a | 2023-08-22 | RSD | 7.19% | 1.704% | 1.055% | <LOD | <LOD | 1.477% | 2.839% | 7.87% | 39.51% | 23.4% | 7.26% | 4.817% | 21.42% | <LOD | 5.04% | 5.267% | 13.95% | 46.63% | 2.316% | 5.005% | 10.55% | 8.685% | 15.27% | 44.95% | 61.5% | 21.53% | 38.31% | 26.98% | 51.93% | 55.16% | <LOD | 8.43% | 3.447% | <LOD | 57.64% | <LOD | 0.6878%
7 [Col] | 2023-08-22 | NIST 679 | standard | rep 1 | <LOD | 91548.2 | 220027.9 | <LOD | <LOD | 24321 | 1647.3 | 5430.7 | 228.1 | 92 | 1706.9 | 92568.5 | <LOD | 61.2 | 65.8 | 116.2 | 22.5 | 10.2 | 188 | 53.6 | 40.2 | 152.1 | 19.7 | 14.3 | <LOD | 35.4 | 38.3 | 79.4 | 925.4 | 37 | <LOD | 8.3 | 15.9 | <LOD | 8.4 | 4.2 | 704838.9
8 [Col] | 2023-08-22 | NIST 679 | standard | rep 2 | <LOD | 90270.6 | 220672 | <LOD | <LOD | 24199.8 | 1641.3 | 5421.8 | 100.1 | 34.2 | 1702.4 | 92502.9 | 168.3 | 80 | 67 | 123.3 | 27.2 | 7.5 | 185.7 | 54.6 | 33.1 | 161.7 | 25.3 | 16.9 | 49.8 | 51.1 | 85.9 | 111.8 | 757.9 | 54.3 | <LOD | <LOD | <LOD | 22 | <LOD | <LOD | 705380.6
9 [Col] | 2023-08-22 | NIST 679 | standard | rep 3 | <LOD | 89989.8 | 220593.2 | <LOD | <LOD | 24447 | 1510.7 | 5801.8 | 190.5 | 32.1 | 1630.9 | 93159.4 | 97 | 70.2 | 90.6 | 122.1 | 23.3 | 6 | 202.1 | 49.5 | 43.4 | 173.6 | 22.9 | 9.7 | 27.7 | 36.8 | 52.5 | 110.9 | <LOD | <LOD | <LOD | <LOD | 20.2 | <LOD | 37.8 | <LOD | 705423.2
10 [Col] | 2023-08-22 | NIST 679 | standard | rep 4 | <LOD | 92589.1 | 221543.6 | <LOD | <LOD | 23967 | 1395.2 | 5071.1 | 234.3 | <LOD | 1792.1 | 90363 | 240.2 | 48.4 | 56.3 | 127.5 | 26.2 | 4.4 | 187.8 | 50.3 | 36.4 | 157.4 | 20.2 | 6.3 | <LOD | <LOD | 44.6 | 47.3 | 847.6 | 20.7 | <LOD | 12.5 | 17.5 | <LOD | 24.1 | <LOD | 705512.5
11 [Col] | 2023-08-22 | NIST 679 | standard | rep 5 | <LOD | 86753.5 | 214972.9 | <LOD | <LOD | 23411.9 | 1555.3 | 5461.1 | 280.1 | 87.4 | 1534.2 | 90771.8 | 264.1 | 61.9 | 55.3 | 114.6 | 16.3 | 3.2 | 188.1 | 53.7 | 37.4 | 166.4 | 19.7 | 9.3 | <LOD | 23.6 | 43.6 | 105.5 | <LOD | <LOD | <LOD | 5.3 | 21.9 | <LOD | 9.1 | 5.4 | 716433.3
|  | NIST 679 | 2023-08-22 | average | <LOD | 90230.24 | 219561.92 | <LOD | <LOD | 24069.34 | 1549.96 | 5437.3 | 206.62 | 61.425 | 1673.3 | 91873.12 | 192.4 | 64.34 | 67 | 120.74 | 23.1 | 6.26 | 190.34 | 52.34 | 38.1 | 162.24 | 21.56 | 11.3 | <LOD | 36.725 | 52.98 | 90.98 | 843.6333 | 37.3333 | <LOD | 8.7 | 18.875 | <LOD | 19.85 | <LOD | 707517.7
|  | NIST 679 | 2023-08-22 | stdev | <LOD | 2205.017 | 2621.9799 | <LOD | <LOD | 407.9462 | 104.0896 | 258.7532 | 67.5153 | 32.7144 | 96.4829 | 1227.6055 | 75.5153 | 11.7204 | 14.2283 | 5.3012 | 4.2738 | 2.7364 | 6.6478 | 2.2788 | 3.9013 | 8.2597 | 2.4775 | 4.2403 | <LOD | 11.2645 | 19.0897 | 27.7631 | 83.8204 | 16.8025 | <LOD | 3.6166 | 2.6862 | <LOD | 13.9873 | <LOD | 4990.963
|  | NIST 679 | 2023-08-22 | RSD | <LOD | 2.444% | 1.194% | <LOD | <LOD | 1.695% | 6.716% | 4.759% | 32.68% | 53.26% | 5.766% | 1.336% | 39.25% | 18.22% | 21.24% | 4.391% | 18.5% | 43.71% | 3.493% | 4.354% | 10.24% | 5.091% | 11.49% | 37.52% | <LOD | 30.67% | 36.03% | 30.52% | 9.936% | 45.01% | <LOD | 41.57% | 14.23% | <LOD | 70.46% | <LOD | 0.7054%
12 [Col] | 2023-08-22 | SARM 69 | standard | rep 1 | 12959.4 | 78232.8 | 268080.7 | 152.8 | <LOD | 17555.2 | 14857.7 | 4707.2 | 126.4 | 152.1 | 1050.6 | 51302.7 | 173.6 | 71 | 54.6 | 62 | 7 | 1.5 | 78.1 | 104.8 | 28.6 | 251.1 | 9.7 | 9.5 | <LOD | 33.6 | 43.2 | 44.4 | 999.8 | <LOD | 7.6 | 9 | 21.9 | <LOD | <LOD | <LOD | 705621.3
13 [Col] | 2023-08-22 | SARM 69 | standard | rep 2 | 12577.8 | 77957.6 | 268298.5 | 151.8 | <LOD | 17717 | 14568.4 | 4505.2 | 215.5 | 262.1 | 963.2 | 50566.8 | 235 | 59.1 | 82.1 | 71.8 | 2.8 | 3 | 73.6 | 103.8 | 26.4 | 284.6 | 5.7 | <LOD | 49.7 | 45.8 | 30.3 | 56.9 | 1201.9 | <LOD | 3.8 | <LOD | 19.1 | <LOD | 7.3 | <LOD | 706662.1
14 [Col] | 2023-08-22 | SARM 69 | standard | rep 3 | <LOD | 79959.3 | 272471.9 | 188.9 | <LOD | 18405.7 | 14985.1 | 4337.9 | 327.4 | 177.1 | 940.2 | 52366.2 | <LOD | 66.9 | 102 | 71.7 | 10.1 | 2.9 | 76.7 | 97 | 27.1 | 307 | 5.1 | 10.1 | 18.9 | 20 | 36.2 | 54.6 | 635.2 | 15.7 | <LOD | <LOD | 17.6 | <LOD | 18.1 | 7.5 | 702944.3
15 [Col] | 2023-08-22 | SARM 69 | standard | rep 4 | 11559.5 | 78728.9 | 270051 | 200.3 | <LOD | 17752.4 | 15285.8 | 4774 | 343.3 | 103.2 | 904.4 | 51802.8 | 241.7 | 39.5 | 84.6 | 70.9 | 12.7 | 3.7 | 75.3 | 101.8 | 40.4 | 263.5 | <LOD | 10.2 | 17.8 | 32.4 | 55.5 | 91.3 | 429.2 | <LOD | <LOD | <LOD | <LOD | 16.8 | <LOD | <LOD | 704430.7
16 [Col] | 2023-08-22 | SARM 69 | standard | rep 5 | 13470.7 | 79686.8 | 275295.4 | 235.1 | <LOD | 18389.5 | 15353.8 | 5409.1 | <LOD | 194.7 | 1005.6 | 52765.9 | <LOD | 44.1 | 137.8 | 73.2 | 10.4 | <LOD | 76.6 | 105.6 | 33.4 | 306.2 | 11.5 | 11.4 | <LOD | 50.8 | 63.4 | 65.8 | 905.6 | <LOD | <LOD | <LOD | 16.3 | <LOD | 27.7 | <LOD | 694334.4
|  | SARM 69 | 2023-08-22 | average | 12641.85 | 78913.08 | 270839.5 | 185.78 | <LOD | 17963.96 | 15010.16 | 4746.68 | 253.15 | 177.84 | 972.8 | 51760.88 | 216.7667 | 56.12 | 92.22 | 69.92 | 8.6 | 2.775 | 76.06 | 102.6 | 31.18 | 282.48 | 8 | 10.3 | 28.8 | 36.52 | 45.72 | 62.6 | 834.34 | <LOD | <LOD | <LOD | 18.725 | <LOD | 17.7 | <LOD | 702798.56
|  | SARM 69 | 2023-08-22 | stdev | 808.9941 | 880.7474 | 3049.9658 | 34.9834 | <LOD | 402.8182 | 321.3603 | 408.2031 | 101.8547 | 58.3486 | 56.9257 | 868.0424 | 37.5332 | 13.8493 | 30.6142 | 4.5041 | 3.8243 | 0.9215 | 1.695 | 3.438 | 5.8328 | 25.0665 | 3.1005 | 0.79582 | 18.1083 | 12.129 | 13.6289 | 17.7599 | 304.6927 | <LOD | <LOD | <LOD | 2.4061 | <LOD | 10.2059 | <LOD | 4930.059
|  | SARM 69 | 2023-08-22 | RSD | 6.399% | 1.116% | 1.126% | 18.83% | <LOD | 2.242% | 2.141% | 8.6% | 40.23% | 32.81% | 5.852% | 1.677% | 17.32% | 24.68% | 33.2% | 6.442% | 44.47% | 33.21% | 2.228% | 3.351% | 18.71% | 8.874% | 38.76% | 7.726% | 62.88% | 33.21% | 29.81% | 28.37% | 36.52% | <LOD | <LOD | <LOD | 12.85% | <LOD | 57.66% | <LOD | 0.7015%
1 [Col] | 2024-05-14 | NIST 2711a | std | rep 1 | 14782 | 69584 | 265770 | <LOD | <LOD | 26958 | 24551 | 3181 | <LOD | 87 | 655 | 29057 | 50 | 17 | 163 | 363 | 132 | 5 | 119 | 251 | 31 | 351 | 25 | 7 | 12 | 80 | 23 | 82 | 1518 | 38 | <LOD | <LOD | 1319 | <LOD | <LOD | <LOD | 722799
2 [Col] | 2024-05-14 | NIST 2711a | std | rep 2 | 13594 | 69569 | 260279 | <LOD | <LOD | 26870 | 24961 | <LOD | 71 | 104 | 690 | 28836 | 92 | 21 | 178 | 398 | 153 | 3 | 124 | 252 | 28 | 303 | 18 | 10 | 17 | 67 | 39 | 87 | 1535 | 41 | <LOD | 10 | 1352 | <LOD | <LOD | <LOD | 729210
3 [Col] | 2024-05-14 | NIST 2711a | std | rep 3 | 12937 | 69986 | 266201 | <LOD | <LOD | 26372 | 24421 | <LOD | 126 | 97 | 613 | 27918 | 165 | 22 | 145 | 400 | 130 | 3 | 120 | 250 | 29 | 318 | 26 | 11 | 19 | 66 | 27 | 82 | 1962 | 30 | <LOD | 10 | 1294 | 36 | <LOD | <LOD | 725456
4 [Col] | 2024-05-14 | NIST 2711a | std | rep 4 | 14818 | 69948 | 267316 | <LOD | <LOD | 27258 | 25880 | 3342 | <LOD | 99 | 790 | 29792 | <LOD | 33 | 177 | 418 | 143 | <LOD | 121 | 267 | 29 | 320 | 14 | 5 | 16 | 87 | 86 | 95 | 1510 | <LOD | <LOD | 9 | 1379 | <LOD | 9 | 5 | 718820
5 [Col] | 2024-05-14 | NIST 2711a | std | rep 5 | 13830 | 70391 | 268520 | <LOD | <LOD | 26742 | 25026 | 3436 | 203 | <LOD | 608 | 29764 | <LOD | 38 | 177 | 388 | 123 | <LOD | 120 | 257 | 26 | 272 | 23 | 13 | <LOD | 79 | 31 | 80 | 1418 | <LOD | <LOD | <LOD | 1325 | <LOD | 7 | <LOD | 719716
|  | NIST 2711a | 2024-05-14 | average | 13992.2 | 69895.6 | 265617.2 | <LOD | <LOD | 26840 | 24967.8 | 3319.6667 | 133.3333 | 96.75 | 671.2 | 29073.4 | 102.3333 | 26.2 | 168 | 393.4 | 136.2 | 3.6667 | 120.8 | 255.4 | 28.6 | 312.8 | 21.2 | 9.2 | 16 | 75.8 | 41.2 | 85.2 | 1588.6 | 36.3333 | <LOD | 9.6667 | 1333.8 | <LOD | <LOD | <LOD | 723200.2
|  | NIST 2711a | 2024-05-14 | stdev | 806.8533 | 339.1538 | 3169.0976 | <LOD | <LOD | 323.3326 | 571.8948 | 128.9586 | 66.3049 | 7.1356 | 74.3485 | 772.1546 | 58.1922 | 8.8713 | 14.2829 | 20.1445 | 11.8195 | 1.1547 | 1.9235 | 7.0214 | 1.8166 | 28.7176 | 5.0695 | 3.1937 | 2.9439 | 9.0388 | 25.7332 | 6.0581 | 213.6394 | 5.6862 | <LOD | 0.57735 | 32.6144 | <LOD | <LOD | <LOD | 4263.0122
|  | NIST 2711a | 2024-05-14 | RSD | 5.766% | 0.4852% | 1.193% | <LOD | <LOD | 1.205% | 2.291% | 3.885% | 49.73% | 7.375% | 11.08% | 2.656% | 56.87% | 33.86% | 8.502% | 5.121% | 8.678% | 31.49% | 1.592% | 2.749% | 6.352% | 9.181% | 23.91% | 34.71% | 18.4% | 11.92% | 62.46% | 7.11% | 13.45% | 15.65% | <LOD | 5.973% | 2.445% | <LOD | <LOD | <LOD | 0.5895%
6 [Col] | 2024-05-14 | NIST 679 | std | rep 1 | <LOD | 89060 | 218471 | 466 | <LOD | 24860 | 1920 | 5720 | 417 | 88 | 1537 | 88368 | 178 | 51 | 89 | 134 | 25 | 6 | 186 | 49 | 37 | 161 | 15 | <LOD | 15 | 64 | 131 | 80 | <LOD | 30 | 4 | <LOD | <LOD | <LOD | 19 | <LOD | 711778
7 [Col] | 2024-05-14 | NIST 679 | std | rep 2 | <LOD | 88828 | 216347 | 462 | <LOD | 24403 | 1925 | 5176 | 174 | 84 | 1536 | 85986 | 270 | 47 | 62 | 123 | 18 | 4 | 189 | 44 | 35 | 155 | 16 | 11 | 42 | 32 | 43 | 62 | 548 | <LOD | <LOD | <LOD | <LOD | 24 | <LOD | <LOD | 716953
8 [Col] | 2024-05-14 | NIST 679 | std | rep 3 | <LOD | 87222 | 214286 | 430 | <LOD | 24395 | 1930 | 5333 | 244 | 95 | 1658 | 87508 | 225 | 41 | 63 | 124 | 12 | <LOD | 190 | 48 | 40 | 165 | 18 | 10 | <LOD | <LOD | 72 | 74 | <LOD | <LOD | <LOD | <LOD | 21 | <LOD | 17 | <LOD | 718860
9 [Col] | 2024-05-14 | NIST 679 | std | rep 4 | <LOD | 86831 | 214539 | 418 | <LOD | 24004 | 1894 | 5361 | 193 | 55 | 1696 | 86980 | 299 | 50 | 49 | 135 | 25 | 6 | 189 | 49 | 37 | 151 | 15 | <LOD | 43 | 49 | 53 | 116 | 1348 | <LOD | <LOD | <LOD | <LOD | <LOD | 10 | <LOD | 718231
10 [Col] | 2024-05-14 | NIST 679 | std | rep 5 | <LOD | 86280 | 213695 | 427 | <LOD | 23776 | 1941 | 4949 | 244 | 92 | 1579 | 85726 | 302 | 48 | 69 | 114 | 14 | 4 | 176 | 42 | 46 | 159 | 19 | 11 | <LOD | 24 | 83 | 46 | 886 | <LOD | <LOD | <LOD | <LOD | <LOD | 12 | <LOD | 721842
|  | NIST 679 | 2024-05-14 | average | <LOD | 87644.2 | 215467.6 | 440.6 | <LOD | 24287.6 | 1922 | 5307.8 | 254.4 | 82.8 | 1601.2 | 86913.6 | 254.8 | 47.4 | 66.4 | 126 | 18.8 | 5 | 186 | 46.4 | 39 | 158.2 | 16.6 | 10.6667 | 33.3333 | 42.25 | 76.4 | 75.6 | 927.3333 | <LOD | <LOD | <LOD | <LOD | <LOD | 14.5 | <LOD | 717532.8
|  | NIST 679 | 2024-05-14 | stdev | <LOD | 1235.5627 | 1948.9961 | 21.8586 | <LOD | 416.7305 | 17.4786 | 282.614 | 96.0328 | 16.0842 | 72.5927 | 1089.0174 | 52.9122 | 3.9115 | 14.5877 | 8.6891 | 6.0581 | 1.1547 | 5.7879 | 3.2094 | 4.3012 | 5.4037 | 1.8166 | 0.57735 | 15.885 | 17.8582 | 34.3045 | 26.0538 | 401.5985 | <LOD | <LOD | <LOD | <LOD | <LOD | 4.2032 | <LOD | 3683.4779
|  | NIST 679 | 2024-05-14 | RSD | <LOD | 1.41% | 0.9045% | 4.961% | <LOD | 1.716% | 0.9094% | 5.325% | 37.75% | 19.43% | 4.534% | 1.253% | 20.77% | 8.252% | 21.97% | 6.896% | 32.22% | 23.09% | 3.112% | 6.917% | 11.03% | 3.416% | 10.94% | 5.413% | 47.66% | 42.27% | 44.9% | 34.46% | 43.31% | <LOD | <LOD | <LOD | <LOD | <LOD | 28.99% | <LOD | 0.5134%
11 [Col] | 2024-05-14 | SARM 69 | std | rep 1 | <LOD | 71191 | 250042 | 854 | <LOD | 17196 | 15037 | 4521 | 108 | 205 | 857 | 49523 | <LOD | 60 | 62 | 69 | 10 | 4 | 77 | 95 | 36 | 320 | 11 | 6 | 16 | 32 | 42 | 49 | 597 | <LOD | <LOD | <LOD | <LOD | <LOD | 7 | <LOD | 733530
12 [Col] | 2024-05-14 | SARM 69 | std | rep 2 | <LOD | 70552 | 248683 | 821 | <LOD | 17164 | 14365 | 4602 | <LOD | 195 | 929 | 48913 | <LOD | 57 | 76 | 65 | 5 | 3 | 73 | 88 | 30 | 279 | 7 | 9 | <LOD | <LOD | 58 | 69 | 1064 | 20 | <LOD | <LOD | <LOD | 23 | <LOD | 4 | 735887
13 [Col] | 2024-05-14 | SARM 69 | std | rep 3 | 13439 | 71516 | 257006 | 882 | <LOD | 17682 | 15379 | 4552 | 168 | 152 | 885 | 50340 | 126 | 56 | 68 | 72 | 8 | 3 | 78 | 111 | 31 | 330 | 8 | 10 | 31 | 31 | 38 | 56 | 686 | 25 | <LOD | <LOD | 18 | <LOD | 9 | <LOD | 720882
14 [Col] | 2024-05-14 | SARM 69 | std | rep 4 | <LOD | 69959 | 252919 | 865 | <LOD | 17135 | 15238 | 4527 | 202 | 161 | 904 | 49328 | 228 | 35 | 91 | 71 | 7 | 3 | 78 | 97 | 32 | 303 | 10 | 14 | <LOD | 36 | 70 | 87 | <LOD | 25 | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | 732135
15 [Col] | 2024-05-14 | SARM 69 | std | rep 5 | <LOD | 69631 | 253411 | 786 | <LOD | 17271 | 14651 | 4590 | 240 | 156 | 873 | 49049 | 184 | 53 | 73 | 75 | 5 | 6 | 71 | 91 | 30 | 264 | 8 | 8 | 14 | 34 | 57 | 69 | 621 | 33 | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | 732008
|  | SARM 69 | 2024-05-14 | average | <LOD | 70569.8 | 252412.2 | 841.6 | <LOD | 17289.6 | 14934 | 4558.4 | 179.5 | 173.8 | 889.6 | 49430.6 | 179.3333 | 52.2 | 74 | 70.4 | 7 | 3.8 | 75.4 | 96.4 | 31.8 | 299.2 | 8.8 | 9.4 | 20.3333 | 33.25 | 53 | 66 | 742 | 25.75 | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | 730888.4
|  | SARM 69 | 2024-05-14 | stdev | <LOD | 796.229 | 3235.824 | 38.24 | <LOD | 225.1451 | 419.7916 | 36.487 | 56.0089 | 24.3865 | 27.9249 | 561.1598 | 51.1599 | 9.9348 | 10.8858 | 3.7148 | 2.1213 | 1.3038 | 3.2094 | 8.8769 | 2.49 | 27.5808 | 1.6432 | 2.9665 | 9.2916 | 2.2174 | 13 | 14.5602 | 217.9342 | 5.3774 | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | 5807.3146
|  | SARM 69 | 2024-05-14 | RSD | <LOD | 1.128% | 1.282% | 4.544% | <LOD | 1.302% | 2.811% | 0.8004% | 31.2% | 14.03% | 3.139% | 1.135% | 28.53% | 19.03% | 14.71% | 5.277% | 30.3% | 34.31% | 4.256% | 9.208% | 7.83% | 9.218% | 18.67% | 31.56% | 45.7% | 6.669% | 24.53% | 22.06% | 29.37% | 20.88% | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | 0.7946%
Formulas:
component = [Date, row above]
component = [component, row above]

## % error
sample | Date | component | Mg | Al | Si | K | Ca | Ti | Mn | Fe | Ni | Zn | Rb | Sr | Y | Zr | Nb | Pb
NIST 2711a |  | certified value | 10700 | 67200 | 314000 | 25300 | 24200 | 3170 | 675 | 28200 | 21.7 | 414 |  | 242 |  |  |  | 1400
NIST 679 |  | certified value | Mg: 7552; Al: 110100; Si: 243400; K: 24330; Ca: 1628; Ti: 5770; Fe: 90500; Sr: 73.4
SARM 69 |  | certified value | 11200 | 76200 | 311000 | 16300 | 16900 | 4660 | 999 | 50200 | 53 | 68
NIST 2711a | 2023-06-06 | average | 15578.8 | 68412.2 | 260074.6 | 25260.4 | 22645.6 | 3418.6667 | 678 | 29323.8 | 31.8 | 377.4 | 121.4 | 261.8 | 32 | 290.8 | 19 | 1368.8
NIST 679 | 2023-06-06 | average | 11664 | 89875.2 | 216946.2 | 23425.4 | 1438.6 | 5218.2 | 1457 | 90143 | 58.8 | 111.6 | 189.2 | 49.4 | 43.2 | 160.2 | 19.2 | 17.6667
SARM 69 | 2023-06-06 | average | 13882 | 77194.8 | 259341.2 | 17268.6 | 14056.8 | 4696.8 | 928.6 | 50867.6 | 75.6 | 66.4 | 73.2 | 104.2 | 30.6 | 289.6 | 11.25 | 19
NIST 2711a | 2023-06-07 | average | 14937.8 | 68794.8 | 259959.4 | 25310.2 | 22544.4 | 3339.4 | 723.2 | 29033.2 | 33.2 | 390 | 122.6 | 257.6 | 31 | 317.2 | 21 | 1348.4
NIST 679 | 2023-06-07 | average | 11502.6 | 90394.6 | 217698.2 | 23352.6 | 1419.6 | 5198.4 | 1575.8 | 90015 | 54.4 | 125.6 | 186 | 53.2 | 38.8 | 161.2 | 21.2 | <LOD
SARM 69 | 2023-06-07 | average | 13061.4 | 77698.4 | 263483.8 | 17372.6 | 14074 | 4487.4 | 965.6 | 51190.8 | 59.8 | 63 | 75.8 | 101 | 27.6 | 288.2 | 10.4 | <LOD
NIST 2711a | 2023-06-08 | average | 14911 | 67805.8 | 258990.2 | 25023.4 | 22571.8 | 3281.6 | 652.4 | 28777.2 | 34 | 394.8 | 119 | 256.4 | 32.2 | 305.2 | 19.8 | 1345.2
NIST 679 | 2023-06-08 | average | 11883.25 | 93048 | 221805.2 | 23922.2 | 1392.6 | 5290.8 | 1626.8 | 91304.8 | 60 | 119 | 189 | 48.4 | 35.8 | 162 | 18.6 | <LOD
SARM 69 | 2023-06-08 | average | 14003 | 77945.2 | 264277.8 | 17472.8 | 13983.4 | 4549.8 | 977.6 | 51416.8 | 53.8 | 65.2 | 78.2 | 104.8 | 28.8 | 284.8 | 9.4 | <LOD
NIST 2711a | 2023-06-09 | average | 15003.2 | 69240.2 | 259289.4 | 25062.8 | 22565 | 3418.75 | 667.4 | 28790 | 28.4 | 393.4 | 119.8 | 260 | 32.8 | 306.2 | 18 | 1340.4
NIST 679 | 2023-06-09 | average | 11692 | 91550.6 | 219199.2 | 23586 | 1425.2 | 5263.8 | 1595.2 | 89638.4 | 53.6 | 127 | 185 | 49.8 | 39.2 | 158.8 | 24.2 | <LOD
SARM 69 | 2023-06-09 | average | 12889.2 | 77714 | 262761.4 | 17525.2 | 13988.2 | 4951.2 | 883.2 | 51602.8 | 69.6 | 66.6 | 77.2 | 106.2 | 34.2 | 316.8 | 9.6 | 20
NIST 2711a | 2023-06-13 | average | 15753.6 | 70855 | 265700.6 | 25599.6 | 23029 | 3510.75 | 654.4 | 28756 | 27.2 | 387.4 | 120.6 | 254.2 | 30.2 | 300.2 | 19.2 | 1324.8
NIST 679 | 2023-06-13 | average | 11909 | 91530.6 | 220261.4 | 23823.8 | 1480.2 | 5438 | 1677.4 | 91333.8 | 60.8 | 120 | 186 | 47.8 | 42.4 | 171.4 | 20.2 | <LOD
SARM 69 | 2023-06-13 | average | 13505 | 78044.8 | 265696 | 17300.2 | 14366.6 | 4376.8 | 1000.4 | 51333.8 | 59.8 | 66.8 | 72.6 | 99.2 | 31 | 297 | 10.75 | <LOD
NIST 2711a | 2023-06-14 | average | 14670 | 68978.4 | 261282.4 | 25298.4 | 22848.4 | 3370.25 | 644.6 | 28958.4 | 25.75 | 388.6 | 119.2 | 258.8 | 34 | 294 | 18.6 | 1347.4
NIST 679 | 2023-06-14 | average | 11518.5 | 92470 | 220459.4 | 23815.8 | 1411.2 | 5371 | 1571.4 | 92211 | 67 | 124 | 189.4 | 48.4 | 40.4 | 162.8 | 23.6 | <LOD
SARM 69 | 2023-06-14 | average | 12990 | 77269.8 | 262662.6 | 17398 | 14041.2 | 4380.4 | 967.2 | 51036.4 | 64.2 | 67.2 | 79.4 | 105.6 | 33.4 | 286.4 | 10.5 | <LOD
NIST 2711a | 2023-06-15 | average | 14928.8 | 69620 | 264482.8 | 25575.2 | 22913.4 | 3364.75 | 719.8 | 29251.2 | 30.2 | 386.8 | 121.4 | 262.2 | 32 | 309.2 | 18.6 | 1358.6
NIST 679 | 2023-06-15 | average | 11619 | 91659.6 | 219834 | 23862.4 | 1440.2 | 5643.2 | 1634.2 | 91949.6 | 63.2 | 120 | 195 | 54.4 | 42.2 | 163.8 | 25.8 | <LOD
SARM 69 | 2023-06-15 | average | 13325 | 78403.6 | 266428.8 | 17665.4 | 14177.4 | 4502.8 | 1001.6 | 51298.2 | 57.2 | 68.8 | 78 | 107.6 | 33.8 | 294.2 | 9.25 | <LOD
NIST 2711a | 2023-06-16 | average | 14881.2 | 68553 | 261041.4 | 25334.8 | 22597 | 3569 | 648.2 | 29185.8 | 27.8 | 390.4 | 121.6 | 261.4 | 32.4 | 311.8 | 21.6 | 1345.2
NIST 679 | 2023-06-16 | average | 12144.5 | 93018.8 | 221863.2 | 24109.4 | 1463.4 | 5635.8 | 1605.4 | 92454.4 | 60.4 | 134 | 191.2 | 51.2 | 36 | 165.8 | 23.8 | <LOD
SARM 69 | 2023-06-16 | average | 12661.8 | 78923 | 267957.2 | 17574.4 | 14322.4 | 4661.2 | 966 | 51047.6 | 59.6 | 72.4 | 77 | 106 | 31.4 | 300 | 9.4 | 18.3333
NIST 2711a | 2023-07-18 | average | 11929.0667 | 66025.64 | 256225.84 | 24470.8 | 22160.94 | 3405.3667 | 646 | 28188.76 | 23.775 | 379.24 | 116.5 | 256.42 | 32.18 | 296.58 | 20.44 | 1302.76
NIST 679 | 2023-07-18 | average | <LOD | 85125.42 | 209829.22 | 21927.9 | 1499.025 | 4837.6 | 1474.9 | 82939.26 | 33.82 | 113.38 | 174.48 | 46.075 | 37.04 | 147.42 | 22.78 | <LOD
SARM 69 | 2023-07-18 | average | 11898.78 | 77017.76 | 265526.98 | 17140.68 | 14026.78 | 4582.5 | 928.88 | 50984.38 | 51.58 | 61.92 | 73.06 | 99.6 | 31.4 | 305.06 | 10.325 | <LOD
NIST 2711a | 2023-07-28 | average | 13401.38 | 68875.7 | 264999.7 | 25785.08 | 23222.42 | 3519.18 | 679.46 | 28993.22 | 23.175 | 391.16 | 123.5 | 263.88 | 31.14 | 288.26 | 21.66 | 1356.42
NIST 679 | 2023-07-28 | average | <LOD | 86841.46 | 213132.88 | 22953.36 | 1426.2 | 5496.62 | 1580.9 | 88192.48 | 53.88 | 122.16 | 179.64 | 44.26 | 39.12 | 152.84 | 18.38 | <LOD
SARM 69 | 2023-07-28 | average | 13179.4 | 78825.82 | 272273.7 | 17672.12 | 14695.02 | 4870.68 | 935.4 | 51227.6 | 51.8 | 68.7 | 77.12 | 102.26 | 30.66 | 279.82 | 7.62 | 17.9667
NIST 2711a | 2023-07-31 | average | 13921.275 | 70366.48 | 268191.98 | 26021.34 | 23788.14 | 3330.88 | 646.76 | 29178.62 | 22.85 | 395.82 | 120.98 | 258.54 | 30.98 | 298.58 | 22.2 | 1349.84
NIST 679 | 2023-07-31 | average | <LOD | 88741.6 | 217822.1 | 23655 | 1501.76 | 5278.06 | 1619.08 | 90705.5 | 37.74 | 116.16 | 184.26 | 48.76 | 39.7 | 156.3 | 20.92 | <LOD
SARM 69 | 2023-07-31 | average | 12216.8667 | 73715.24 | 254816.04 | 16073.58 | 13271.22 | 4461.625 | 862.78 | 44797.86 | 52.7 | 59.56 | 65.96 | 103.025 | 26.4 | 234.9 | 10.175 | <LOD
NIST 2711a | 2023-08-01 | average | 13252.825 | 70567.2 | 266270.32 | 25939.36 | 23526.24 | 3600.15 | 674.36 | 28700.74 | 23.74 | 406.44 | 119.44 | 259.76 | 36.18 | 314.4 | 18.82 | 1327.74
NIST 679 | 2023-08-01 | average | <LOD | 92070.52 | 221329.94 | 24111.8 | 1537.12 | 5279.04 | 1628.04 | 91428.32 | 55.52 | 125.06 | 188.36 | 48.1 | 37.4 | 157.48 | 14.54 | <LOD
SARM 69 | 2023-08-01 | average | 12349.08 | 78116.32 | 268881.54 | 17558.38 | 14462.64 | 4694.02 | 957.52 | 51647.16 | 51.68 | 61.84 | 78.18 | 105.34 | 34.32 | 288.28 | 11.025 | <LOD
NIST 2711a | 2023-08-02 | average | 14180.14 | 70385.88 | 268553.24 | 25989.78 | 23452.46 | 3754.3 | 637.88 | 29371.34 | 19.05 | 400.52 | 116.84 | 266.64 | 33.5 | 299.62 | 20.98 | 1321.38
NIST 679 | 2023-08-02 | average | <LOD | 92953.3 | 223590.88 | 24542.12 | 1521.08 | 5482.02 | 1631.64 | 91793.66 | 51.76 | 126.8 | 190.56 | 52.2 | 37.58 | 161.24 | 20.96 | <LOD
SARM 69 | 2023-08-02 | average | 12331.56 | 79077.86 | 272786.2 | 17644.28 | 15054.48 | 4718.86 | 961.56 | 51028.72 | 64.18 | 68.12 | 74.24 | 102.08 | 32.2 | 282.28 | 8.92 | 17.1
NIST 2711a | 2023-08-03 | average | 13857.16 | 71713.52 | 271186.02 | 26331.82 | 24181.82 | 3301.275 | 696.22 | 29164.54 | 32.625 | 393.58 | 118.66 | 265.78 | 36.34 | 300.86 | 17.76 | 1341.84
NIST 679 | 2023-08-03 | average | <LOD | 93929.3 | 225032.26 | 24718.66 | 1585.16 | 5693.12 | 1630.04 | 93143.54 | 46.62 | 125.44 | 192.18 | 51.94 | 40.34 | 160.28 | 17 | <LOD
SARM 69 | 2023-08-03 | average | 12534.08 | 80026.24 | 273109.84 | 17841.38 | 14897.36 | 4762.52 | 926.38 | 51342.04 | 61.28 | 75.3 | 75.96 | 98.74 | 28.8 | 302.3 | 10.28 | 18.1
NIST2711a | 2023-08-22 | average | 14155.78 | 70472.36 | 269331.16 | 26468.22 | 24265.36 | 3272.8 | 637.68 | 29700.6 | <LOD | 409.8 | 120.86 | 269.78 | 33.82 | 289.68 | 22.2 | 1362.86
NIST 679 | 2023-08-22 | average | <LOD | 90230.24 | 219561.92 | 24069.34 | 1549.96 | 5437.3 | 1673.3 | 91873.12 | 64.34 | 120.74 | 190.34 | 52.34 | 38.1 | 162.24 | 21.56 | 18.875
SARM 69 | 2023-08-22 | average | 12641.85 | 78913.08 | 270839.5 | 17963.96 | 15010.16 | 4746.68 | 972.8 | 51760.88 | 56.12 | 69.92 | 76.06 | 102.6 | 31.18 | 282.48 | 8 | 18.725
NIST 2711a | 2024-05-14 | average | 13992.2 | 69895.6 | 265617.2 | 26840 | 24967.8 | 3319.6667 | 671.2 | 29073.4 | 26.2 | 393.4 | 120.8 | 255.4 | 28.6 | 312.8 | 21.2 | 1333.8
NIST 679 | 2024-05-14 | average | <LOD | 87644.2 | 215467.6 | 24287.6 | 1922 | 5307.8 | 1601.2 | 86913.6 | 47.4 | 126 | 186 | 46.4 | 39 | 158.2 | 16.6 | <LOD
SARM 69 | 2024-05-14 | average | <LOD | 70569.8 | 252412.2 | 17289.6 | 14934 | 4558.4 | 889.6 | 49430.6 | 52.2 | 70.4 | 75.4 | 96.4 | 31.8 | 299.2 | 8.8 | <LOD
NIST 2711a | 2023-06-06 | % error | 45.6% | 1.804% | 17.17% | 0.1565% | 6.423% | 7.844% | 0.4444% | 3.985% | 46.54% | 8.841% |  | 8.182% |  |  |  | 2.229%
NIST 679 | 2023-06-06 | % error | Mg: 54.45%; Al: 18.37%; Si: 10.87%; K: 3.718%; Ca: 11.63%; Ti: 9.563%; Fe: 0.3945%; Sr: 32.7%
SARM 69 | 2023-06-06 | % error | 23.95% | 1.306% | 16.61% | 5.942% | 16.82% | 0.7897% | 7.047% | 1.33% | 42.64% | 2.353%
NIST 2711a | 2023-06-07 | % error | 39.61% | 2.373% | 17.21% | 0.04032% | 6.841% | 5.344% | 7.141% | 2.955% | 53% | 5.797% |  | 6.446% |  |  |  | 3.686%
NIST 679 | 2023-06-07 | % error | Mg: 52.31%; Al: 17.9%; Si: 10.56%; K: 4.017%; Ca: 12.8%; Ti: 9.906%; Fe: 0.5359%; Sr: 27.52%
SARM 69 | 2023-06-07 | % error | 16.62% | 1.966% | 15.28% | 6.58% | 16.72% | 3.704% | 3.343% | 1.974% | 12.83% | 7.353%
NIST 2711a | 2023-06-08 | % error | 39.36% | 0.9015% | 17.52% | 1.093% | 6.728% | 3.521% | 3.348% | 2.047% | 56.68% | 4.638% |  | 5.95% |  |  |  | 3.914%
NIST 679 | 2023-06-08 | % error | Mg: 57.35%; Al: 15.49%; Si: 8.872%; K: 1.676%; Ca: 14.46%; Ti: 8.305%; Fe: 0.8893%; Sr: 34.06%
SARM 69 | 2023-06-08 | % error | 25.03% | 2.29% | 15.02% | 7.195% | 17.26% | 2.365% | 2.142% | 2.424% | 1.509% | 4.118%
NIST 2711a | 2023-06-09 | % error | 40.22% | 3.036% | 17.42% | 0.9375% | 6.756% | 7.847% | 1.126% | 2.092% | 30.88% | 4.976% |  | 7.438% |  |  |  | 4.257%
NIST 679 | 2023-06-09 | % error | Mg: 54.82%; Al: 16.85%; Si: 9.943%; K: 3.058%; Ca: 12.46%; Ti: 8.773%; Fe: 0.952%; Sr: 32.15%
SARM 69 | 2023-06-09 | % error | 15.08% | 1.987% | 15.51% | 7.517% | 17.23% | 6.249% | 11.59% | 2.794% | 31.32% | 2.059%
NIST 2711a | 2023-06-13 | % error | 47.23% | 5.439% | 15.38% | 1.184% | 4.839% | 10.75% | 3.052% | 1.972% | 25.35% | 6.425% |  | 5.041% |  |  |  | 5.371%
NIST 679 | 2023-06-13 | % error | Mg: 57.69%; Al: 16.87%; Si: 9.506%; K: 2.081%; Ca: 9.079%; Ti: 5.754%; Fe: 0.9213%; Sr: 34.88%
SARM 69 | 2023-06-13 | % error | 20.58% | 2.421% | 14.57% | 6.136% | 14.99% | 6.077% | 0.1401% | 2.259% | 12.83% | 1.765%
NIST 2711a | 2023-06-14 | % error | 37.1% | 2.646% | 16.79% | 0.006324% | 5.585% | 6.317% | 4.504% | 2.689% | 18.66% | 6.135% |  | 6.942% |  |  |  | 3.757%
NIST 679 | 2023-06-14 | % error | Mg: 52.52%; Al: 16.01%; Si: 9.425%; K: 2.113%; Ca: 13.32%; Ti: 6.915%; Fe: 1.891%; Sr: 34.06%
SARM 69 | 2023-06-14 | % error | 15.98% | 1.404% | 15.54% | 6.736% | 16.92% | 6% | 3.183% | 1.666% | 21.13% | 1.176%
NIST 2711a | 2023-06-15 | % error | 39.52% | 3.601% | 15.77% | 1.088% | 5.317% | 6.144% | 6.637% | 3.728% | 39.17% | 6.57% |  | 8.347% |  |  |  | 2.957%
NIST 679 | 2023-06-15 | % error | Mg: 53.85%; Al: 16.75%; Si: 9.682%; K: 1.922%; Ca: 11.54%; Ti: 2.198%; Fe: 1.602%; Sr: 25.89%
SARM 69 | 2023-06-15 | % error | 18.97% | 2.892% | 14.33% | 8.377% | 16.11% | 3.373% | 0.2603% | 2.188% | 7.925% | 1.176%
NIST 2711a | 2023-06-16 | % error | 39.08% | 2.013% | 16.87% | 0.1375% | 6.624% | 12.59% | 3.97% | 3.496% | 28.11% | 5.7% |  | 8.017% |  |  |  | 3.914%
NIST 679 | 2023-06-16 | % error | Mg: 60.81%; Al: 15.51%; Si: 8.848%; K: 0.9067%; Ca: 10.11%; Ti: 2.326%; Fe: 2.16%; Sr: 30.25%
SARM 69 | 2023-06-16 | % error | 13.05% | 3.573% | 13.84% | 7.818% | 15.25% | 0.02575% | 3.303% | 1.688% | 12.45% | 6.471%
NIST 2711a | 2023-07-18 | % error | 11.49% | 1.748% | 18.4% | 3.277% | 8.426% | 7.425% | 4.296% | 0.03986% | 9.562% | 8.396% |  | 5.959% |  |  |  | 6.946%
NIST 679 | 2023-07-18 | % error | Mg: <LOD; Al: 22.68%; Si: 13.79%; K: 9.873%; Ca: 7.922%; Ti: 16.16%; Fe: 8.354%; Sr: 37.23%
SARM 69 | 2023-07-18 | % error | 6.239% | 1.073% | 14.62% | 5.158% | 17% | 1.663% | 7.019% | 1.563% | 2.679% | 8.941%
NIST 2711a | 2023-07-28 | % error | 25.25% | 2.494% | 15.61% | 1.917% | 4.04% | 11.02% | 0.6607% | 2.813% | 6.797% | 5.517% |  | 9.041% |  |  |  | 3.113%
NIST 679 | 2023-07-28 | % error | Mg: <LOD; Al: 21.12%; Si: 12.44%; K: 5.658%; Ca: 12.4%; Ti: 4.738%; Fe: 2.55%; Sr: 39.7%
SARM 69 | 2023-07-28 | % error | 17.67% | 3.446% | 12.45% | 8.418% | 13.05% | 4.521% | 6.366% | 2.047% | 2.264% | 1.029%
NIST 2711a | 2023-07-31 | % error | 30.11% | 4.712% | 14.59% | 2.851% | 1.702% | 5.075% | 4.184% | 3.47% | 5.3% | 4.391% |  | 6.835% |  |  |  | 3.583%
NIST 679 | 2023-07-31 | % error | Mg: <LOD; Al: 19.4%; Si: 10.51%; K: 2.774%; Ca: 7.754%; Ti: 8.526%; Fe: 0.2271%; Sr: 33.57%
SARM 69 | 2023-07-31 | % error | 9.079% | 3.261% | 18.07% | 1.389% | 21.47% | 4.257% | 13.64% | 10.76% | 0.566% | 12.41%
NIST 2711a | 2023-08-01 | % error | 23.86% | 5.011% | 15.2% | 2.527% | 2.784% | 13.57% | 0.09481% | 1.776% | 9.401% | 1.826% |  | 7.339% |  |  |  | 5.161%
NIST 679 | 2023-08-01 | % error | Mg: <LOD; Al: 16.38%; Si: 9.067%; K: 0.8968%; Ca: 5.582%; Ti: 8.509%; Fe: 1.026%; Sr: 34.47%
SARM 69 | 2023-08-01 | % error | 10.26% | 2.515% | 13.54% | 7.72% | 14.42% | 0.73% | 4.152% | 2.883% | 2.491% | 9.059%
NIST 2711a | 2023-08-02 | % error | 32.52% | 4.741% | 14.47% | 2.726% | 3.089% | 18.43% | 5.499% | 4.154% | 12.21% | 3.256% |  | 10.18% |  |  |  | 5.616%
NIST 679 | 2023-08-02 | % error | Mg: <LOD; Al: 15.57%; Si: 8.139%; K: 0.8718%; Ca: 6.568%; Ti: 4.991%; Fe: 1.429%; Sr: 28.88%
SARM 69 | 2023-08-02 | % error | 10.1% | 3.777% | 12.29% | 8.247% | 10.92% | 1.263% | 3.748% | 1.651% | 21.09% | 0.1765%
NIST 2711a | 2023-08-03 | % error | 29.51% | 6.717% | 13.64% | 4.078% | 0.07512% | 4.141% | 3.144% | 3.42% | 50.35% | 4.932% |  | 9.826% |  |  |  | 4.154%
NIST 679 | 2023-08-03 | % error | Mg: <LOD; Al: 14.69%; Si: 7.546%; K: 1.597%; Ca: 2.631%; Ti: 1.332%; Fe: 2.921%; Sr: 29.24%
SARM 69 | 2023-08-03 | % error | 11.91% | 5.021% | 12.18% | 9.456% | 11.85% | 2.2% | 7.269% | 2.275% | 15.62% | 10.74%
NIST2711a | 2023-08-22 | % error | 32.3% | 4.87% | 14.23% | 4.617% | 0.2701% | 3.243% | 5.529% | 5.321% | <LOD | 1.014% |  | 11.48% |  |  |  | 2.653%
NIST 679 | 2023-08-22 | % error | Mg: <LOD; Al: 18.05%; Si: 9.794%; K: 1.071%; Ca: 4.794%; Ti: 5.766%; Fe: 1.517%; Sr: 28.69%
SARM 69 | 2023-08-22 | % error | 12.87% | 3.56% | 12.91% | 10.21% | 11.18% | 1.86% | 2.623% | 3.109% | 5.887% | 2.824%
NIST 2711a | 2024-05-14 | % error | 30.77% | 4.011% | 15.41% | 6.087% | 3.173% | 4.721% | 0.563% | 3.097% | 20.74% | 4.976% |  | 5.537% |  |  |  | 4.729%
NIST 679 | 2024-05-14 | % error | Mg: <LOD; Al: 20.4%; Si: 11.48%; K: 0.1743%; Ca: 18.06%; Ti: 8.01%; Fe: 3.963%; Sr: 36.78%
SARM 69 | 2024-05-14 | % error | <LOD | 7.389% | 18.84% | 6.071% | 11.63% | 2.18% | 10.95% | 1.533% | 1.509% | 3.529%
average | 2023-06-06 | % error | 41.33% | 7.16% | 14.88% | 3.272% | 11.63% | 6.066% | 3.746% | 1.903% | 44.59% | 5.597% |  | 20.44% |  |  |  | 2.229%
| 2023-06-07 | % error | 36.18% | 7.412% | 14.35% | 3.546% | 12.12% | 6.318% | 5.242% | 1.821% | 32.91% | 6.575% |  | 16.98% |  |  |  | 3.686%
| 2023-06-08 | % error | 40.58% | 6.227% | 13.8% | 3.321% | 12.82% | 4.73% | 2.745% | 1.787% | 29.1% | 4.378% |  | 20.01% |  |  |  | 3.914%
| 2023-06-09 | % error | 36.71% | 7.29% | 14.29% | 3.837% | 12.15% | 7.623% | 6.359% | 1.946% | 31.1% | 3.517% |  | 19.8% |  |  |  | 4.257%
| 2023-06-13 | % error | 41.83% | 8.242% | 13.15% | 3.134% | 9.636% | 7.527% | 1.596% | 1.717% | 19.09% | 4.095% |  | 19.96% |  |  |  | 5.371%
| 2023-06-14 | % error | 35.2% | 6.688% | 13.92% | 2.952% | 11.94% | 6.411% | 3.843% | 2.082% | 19.9% | 3.656% |  | 20.5% |  |  |  | 3.757%
| 2023-06-15 | % error | 37.45% | 7.747% | 13.26% | 3.795% | 10.99% | 3.905% | 3.449% | 2.506% | 23.55% | 3.873% |  | 17.12% |  |  |  | 2.957%
| 2023-06-16 | % error | 37.65% | 7.034% | 13.18% | 2.954% | 10.66% | 4.979% | 3.637% | 2.448% | 20.28% | 6.086% |  | 19.13% |  |  |  | 3.914%
| 2023-07-18 | % error | 8.863% | 8.501% | 15.6% | 6.103% | 11.12% | 8.416% | 5.658% | 3.319% | 6.121% | 8.669% |  | 21.59% |  |  |  | 6.946%
| 2023-07-28 | % error | 21.46% | 9.021% | 13.5% | 5.331% | 9.827% | 6.758% | 3.514% | 2.47% | 4.531% | 3.273% |  | 24.37% |  |  |  | 3.113%
| 2023-07-31 | % error | 19.59% | 9.124% | 14.39% | 2.338% | 10.31% | 5.953% | 8.91% | 4.82% | 2.933% | 8.402% |  | 20.2% |  |  |  | 3.583%
| 2023-08-01 | % error | 17.06% | 7.967% | 12.6% | 3.715% | 7.596% | 7.603% | 2.123% | 1.895% | 5.946% | 5.442% |  | 20.9% |  |  |  | 5.161%
| 2023-08-02 | % error | 21.31% | 8.03% | 11.63% | 3.948% | 6.859% | 8.229% | 4.624% | 2.411% | 16.65% | 1.716% |  | 19.53% |  |  |  | 5.616%
| 2023-08-03 | % error | 20.71% | 8.808% | 11.12% | 5.044% | 4.852% | 2.558% | 5.206% | 2.872% | 32.98% | 7.834% |  | 19.53% |  |  |  | 4.154%
| 2023-08-22 | % error | 22.59% | 8.826% | 12.31% | 5.299% | 5.415% | 3.623% | 4.076% | 3.316% | 5.887% | 1.919% |  | 20.09% |  |  |  | 2.653%
| 2024-05-14 | % error | 30.77% | 10.6% | 15.24% | 4.111% | 10.95% | 4.971% | 5.757% | 2.864% | 11.12% | 4.253% |  | 21.16% |  |  |  | 4.729%

## altitude effect
sample | Date | altitude (m) | component | Mg | Al | Si | K | Ca | Ti | Mn | Fe | Ni | Zn | Rb | Sr | Y | Zr | Nb | Pb | 10% threshold
NIST 2711a |  |  | certified value | 10700 | 67200 | 314000 | 25300 | 24200 | 3170 | 675 | 28200 | 21.7 | 414 |  | 242 |  |  |  | 1400
NIST 679 |  |  | certified value | Mg: 7552; Al: 110100; Si: 243400; K: 24330; Ca: 1628; Ti: 5770; Fe: 90500; Sr: 73.4
SARM 69 |  |  | certified value | 11200 | 76200 | 311000 | 16300 | 16900 | 4660 | 999 | 50200 | 53 | 68
1 | 2023-06-06 | 2610 | ave % error | 41.33% | 7.16% | 14.88% | 3.272% | 11.63% | 6.066% | 3.746% | 1.903% | 44.59% | 5.597% |  | 20.44% |  |  |  | 2.229% | 10%
2 | 2023-06-07 | 2610 | ave % error | 36.18% | 7.412% | 14.35% | 3.546% | 12.12% | 6.318% | 5.242% | 1.821% | 32.91% | 6.575% |  | 16.98% |  |  |  | 3.686% | 10%
3 | 2023-06-08 | 2610 | ave % error | 40.58% | 6.227% | 13.8% | 3.321% | 12.82% | 4.73% | 2.745% | 1.787% | 29.1% | 4.378% |  | 20.01% |  |  |  | 3.914% | 10%
4 | 2023-06-09 | 2610 | ave % error | 36.71% | 7.29% | 14.29% | 3.837% | 12.15% | 7.623% | 6.359% | 1.946% | 31.1% | 3.517% |  | 19.8% |  |  |  | 4.257% | 10%
5 | 2023-06-13 | 1880 | ave % error | 41.83% | 8.242% | 13.15% | 3.134% | 9.636% | 7.527% | 1.596% | 1.717% | 19.09% | 4.095% |  | 19.96% |  |  |  | 5.371% | 10%
6 | 2023-06-14 | 1880 | ave % error | 35.2% | 6.688% | 13.92% | 2.952% | 11.94% | 6.411% | 3.843% | 2.082% | 19.9% | 3.656% |  | 20.5% |  |  |  | 3.757% | 10%
7 | 2023-06-15 | 1880 | ave % error | 37.45% | 7.747% | 13.26% | 3.795% | 10.99% | 3.905% | 3.449% | 2.506% | 23.55% | 3.873% |  | 17.12% |  |  |  | 2.957% | 10%
8 | 2023-06-16 | 1880 | ave % error | 37.65% | 7.034% | 13.18% | 2.954% | 10.66% | 4.979% | 3.637% | 2.448% | 20.28% | 6.086% |  | 19.13% |  |  |  | 3.914% | 10%
9 | 2023-07-18 | 1880 | ave % error | 8.863% | 8.501% | 15.6% | 6.103% | 11.12% | 8.416% | 5.658% | 3.319% | 6.121% | 8.669% |  | 21.59% |  |  |  | 6.946% | 10%
10 | 2023-07-28 | 5 | ave % error | 21.46% | 9.021% | 13.5% | 5.331% | 9.827% | 6.758% | 3.514% | 2.47% | 4.531% | 3.273% |  | 24.37% |  |  |  | 3.113% | 10%
11 | 2023-07-31 | 5 | ave % error | 19.59% | 9.124% | 14.39% | 2.338% | 10.31% | 5.953% | 8.91% | 4.82% | 2.933% | 8.402% |  | 20.2% |  |  |  | 3.583% | 10%
12 | 2023-08-01 | 5 | ave % error | 17.06% | 7.967% | 12.6% | 3.715% | 7.596% | 7.603% | 2.123% | 1.895% | 5.946% | 5.442% |  | 20.9% |  |  |  | 5.161% | 10%
13 | 2023-08-02 | 5 | ave % error | 21.31% | 8.03% | 11.63% | 3.948% | 6.859% | 8.229% | 4.624% | 2.411% | 16.65% | 1.716% |  | 19.53% |  |  |  | 5.616% | 10%
14 | 2023-08-03 | 5 | ave % error | 20.71% | 8.808% | 11.12% | 5.044% | 4.852% | 2.558% | 5.206% | 2.872% | 32.98% | 7.834% |  | 19.53% |  |  |  | 4.154% | 10%
15 | 2023-08-22 | 5 | ave % error | 22.59% | 8.826% | 12.31% | 5.299% | 5.415% | 3.623% | 4.076% | 3.316% | 5.887% | 1.919% |  | 20.09% |  |  |  | 2.653% | 10%
16 | 2024-05-14 | 5 | ave % error | 30.77% | 10.6% | 15.24% | 4.111% | 10.95% | 4.971% | 5.757% | 2.864% | 11.12% | 4.253% |  | 21.16% |  |  |  | 4.729% | 10%
|  |  | R | 0.6899 | -0.75138 | 0.46835 | -0.31496 | 0.7539 | 0.1545 | -0.19209 | -0.54341 | 0.6685 | 0.098583 |  | -0.409 |  |  |  | -0.094088
|  |  | R2 | 0.47597 | 0.56457 | 0.21935 | 0.099198 | 0.56836 | 0.02387 | 0.0369 | 0.29529 | 0.44689 | 0.0097187 |  | 0.16728 |  |  |  | 0.0088525
|  |  | slope | Mg: 6.1547e-05; Al: -7.1065e-06; Ca: 1.5932e-05; Ni: 7.2175e-05
|  |  | intercept | Mg: 0.21685; Al: 0.089251; Ca: 0.079501; Ni: 0.10203

## drift
| NIST 2711a | NIST 2711a average | NIST 2711a average | NIST 2711a average | NIST 2711a average | NIST 2711a average | NIST 2711a average | NIST 2711a average | NIST 2711a average | NIST 2711a average | NIST 2711a average | NIST 2711a average | NIST 2711a average | NIST 2711a average | NIST 2711a average | NIST 2711a average | MO average | UdS average | Cam Average |  | average | average | average | average | average | average | average | average | average | average | average | average | average | average | average |  | average | average | average
| certified | 2023-06-06 | 2023-06-07 | 2023-06-08 | 2023-06-09 | 2023-06-13 | 2023-06-14 | 2023-06-15 | 2023-06-16 | 2023-07-18 | 2023-07-28 | 2023-07-31 | 2023-08-01 | 2023-08-02 | 2023-08-03 | 2023-08-22 |  |  |  |  | 2023-06-06 | 2023-06-07 | 2023-06-08 | 2023-06-09 | 2023-06-13 | 2023-06-14 | 2023-06-15 | 2023-06-16 | 2023-07-18 | 2023-07-28 | 2023-07-31 | 2023-08-01 | 2023-08-02 | 2023-08-03 | 2023-08-22 |  | all | Colombia | Cambridge
|  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 |  |  |  |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15
Mg % | 1.07 | 1.5579 | 1.4938 | 1.4911 | 1.5003 | 1.5754 | 1.467 | 1.4929 | 1.4881 | 1.1065 | 1.3401 | 1.3091 | 1.2556 | 1.418 | 1.3857 | 1.4156 | 1.5108 | 1.5058 | 1.3187 | Mg | 41.33% | 36.18% | 38.71% | 35.01% | 38.06% | 33.66% | 34.18% | 35.55% | 13% | 24.09% | 19.3% | 21.37% | 23.99% | 23.59% | 23.13% | Mg | 29.41% | 36.58% | 21.21%
Al % | 6.72 | 6.8412 | 6.8795 | 6.7806 | 6.924 | 7.0855 | 6.8978 | 6.962 | 6.8553 | 6.6026 | 6.8876 | 7.0366 | 7.0567 | 7.0386 | 7.1714 | 7.0472 | 2610 | 1880 | 6 | Al | 7.16% | 7.412% | 6.227% | 7.29% | 8.242% | 6.688% | 7.747% | 7.034% | 8.501% | 9.021% | 9.124% | 7.967% | 8.03% | 8.808% | 8.826% | Al | 7.872% | 7.225% | 8.611%
Si % | 31.4 | 26.0075 | 25.9959 | 25.899 | 25.9289 | 26.5701 | 26.1282 | 26.4483 | 26.1041 | 25.6226 | 26.5 | 26.8192 | 26.627 | 26.8553 | 27.1186 | 26.9331 |  |  |  | Si | 14.88% | 14.35% | 13.8% | 14.29% | 13.15% | 13.92% | 13.26% | 13.18% | 15.6% | 13.5% | 14.39% | 12.6% | 11.63% | 11.12% | 12.31% | Si | 13.47% | 13.86% | 13.02%
P | 842 | 172.4 | 164.2 | 185.8 | 183.2 | 177.8 | 166.2 | 180.6 | 125 | #DIV/0! | #DIV/0! | 52.2 | 64.2 | 63.9 | #DIV/0! | #DIV/0! |  |  |  | P | 79.52% | 80.5% | 77.93% | 78.24% | 78.88% | 80.26% | 78.55% | 85.15% | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | P
S |  | 227.4 | 155.4 | 181.8 | 185.6 | 170.8 | 157.4 | 131.2 | 145.6 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  |  | K | 3.272% | 3.546% | 3.321% | 3.837% | 3.134% | 2.952% | 3.795% | 2.954% | 6.103% | 5.331% | 2.338% | 3.715% | 3.948% | 5.044% | 5.299% | K | 3.906% | 3.352% | 4.54%
K % | 2.53 | 2.526 | 2.531 | 2.5023 | 2.5063 | 2.56 | 2.5298 | 2.5575 | 2.5335 | 2.4471 | 2.5785 | 2.6021 | 2.5939 | 2.599 | 2.6332 | 2.6468 |  |  |  | Ca | 11.63% | 12.12% | 12.82% | 12.15% | 9.636% | 11.94% | 10.99% | 10.66% | 12.66% | 9.827% | 10.31% | 7.596% | 6.859% | 4.852% | 5.415% | Ca | 9.964% | 11.49% | 8.217%
Ca % | 2.42 | 2.2646 | 2.2544 | 2.2572 | 2.2565 | 2.3029 | 2.2848 | 2.2913 | 2.2597 | 2.2161 | 2.3222 | 2.3788 | 2.3526 | 2.3452 | 2.4182 | 2.4265 |  |  |  | Ti | 3.79% | 6.318% | 4.73% | 6.166% | 6.43% | 5.407% | 3.11% | 4.979% | 5.999% | 6.758% | 9.273% | 5.874% | 8.229% | 1.916% | 3.623% | Ti | 5.507% | 5.116% | 5.953%
Ti | 3170 | 3202.2 | 3339.4 | 3281.6 | 3280.2 | 3406.4 | 3274.8 | 3289.2 | 3569 | 3164.46 | 3519.18 | 3330.88 | 3435.7 | 3754.3 | 3240.28 | 3272.8 |  |  |  | V | 67.7% | 59.23% | 52.66% | 7.187% | 75.96% | 46.84% | 20.2% | 94.96% | 83.93% | 58.61% | 90.3% | 65.61% | 94.24% | 51.9% | 49.69% | V
V | 80.7 | 135.3333 | 128.5 | 123.2 | 86.5 | 142 | 118.5 | 97 | 157.3333 | 148.4333 | 128 | 153.575 | 133.65 | 156.75 | 122.58 | 120.8 |  |  |  | Cr | 16.94% | 30.94% | 36.07% | 29.35% | 32.56% | 46.28% | 29.26% | 27.23% | 35.51% | 14.82% | 25.25% | 36.77% | 18.17% | 12.65% | 24.07% | Cr
Cr | 52.3 | 68.4 | 90 | 100.4 | 94 | 96.5 | 114 | 91.6 | 88.2 | 77.6 | 56.4333 | 57.775 | 80.48 | 49.45 | 49.175 | 48.1333 |  |  |  | Mn | 3.746% | 5.242% | 2.745% | 6.359% | 1.596% | 3.843% | 3.449% | 3.637% | 5.658% | 3.514% | 8.91% | 2.123% | 4.624% | 5.206% | 4.076% | Mn | 4.315% | 3.827% | 4.873%
Mn | 675 | 678 | 723.2 | 652.4 | 667.4 | 654.4 | 644.6 | 719.8 | 648.2 | 646 | 679.46 | 646.76 | 674.36 | 637.88 | 696.22 | 637.68 |  |  |  | Fe | 1.903% | 1.821% | 1.787% | 1.946% | 1.717% | 2.082% | 2.506% | 2.448% | 3.319% | 2.47% | 4.82% | 1.895% | 2.411% | 2.872% | 3.316% | Fe | 2.488% | 2.026% | 3.015%
Fe % | 2.82 | 2.9324 | 2.9033 | 2.8777 | 2.879 | 2.8756 | 2.8958 | 2.9251 | 2.9186 | 2.8189 | 2.8993 | 2.9179 | 2.8701 | 2.9371 | 2.9165 | 2.9701 |  |  |  | Co | 858.2% | 819% | 955.3% | 508.2% | 605.4% | 979.1% | 540.8% | 633% | 950.7% | 754.7% | 820.4% | 713.1% | 804.6% | 696.1% | 969.3% | Co
Co | 9.89 | 135 | 114.25 | 131.5 | 70 | 87.25 | 151.5 | 75.6667 | 75.6667 | 149.15 | 106.08 | 111.3333 | 117.0333 | 100 | 104.56 | 134.9333 |  |  |  | Ni | 44.59% | 32.91% | 29.1% | 31.1% | 19.09% | 12.64% | 23.55% | 20.28% | 1.8% | 4.22% | 14.82% | 5.946% | 20.18% | 23.71% | 12.02% | Ni | 19.73% | 26.66% | 11.81%
Ni | 21.7 | 31.8 | 33.2 | 34 | 28.4 | 27.2 | 22.6 | 30.2 | 27.8 | 21.9 | 20.36 | 18.88 | 23.74 | 17.52 | 28.6 | 17.76 |  |  |  | Cu | 908.6% | 857.1% | 668.6% | 679.2% | 387.8% | 376% | 211% | 243% | 94.13% | 57.22% | 35.07% | 33.94% | 44.76% | 50.68% | 62.1% | Cu
Cu | 140 | 713.2 | 709.4 | 559.8 | 582.8 | 382.8 | 375.2 | 278.6 | 292.8 | 196.78 | 196.18 | 187.26 | 168.68 | 170.44 | 191.08 | 173.2 |  |  |  | Zn | 5.597% | 6.575% | 4.378% | 3.517% | 4.095% | 3.656% | 3.873% | 6.086% | 8.669% | 3.273% | 8.402% | 5.442% | 1.716% | 7.834% | 1.919% | Zn | 5.002% | 4.722% | 5.322%
Zn | 414 | 377.4 | 390 | 394.8 | 393.4 | 387.4 | 388.6 | 386.8 | 390.4 | 379.24 | 391.16 | 395.82 | 406.44 | 400.52 | 393.58 | 409.8 |  |  |  | As | 13.27% | 24.11% | 22.99% | 24.49% | 23.74% | 24.11% | 21.5% | 32.71% | 25.21% | 16.26% | 31.35% | 29.16% | 20.75% | 31.23% | 27.44% | As
As | 107 | 121.2 | 132.8 | 131.6 | 133.2 | 132.4 | 132.8 | 130 | 142 | 133.98 | 124.4 | 140.54 | 138.2 | 129.2 | 140.42 | 136.36 |  |  |  | Sr | 20.44% | 16.98% | 20.01% | 19.8% | 19.96% | 20.5% | 17.12% | 19.13% | 27.01% | 24.37% | 20.2% | 20.9% | 19.53% | 19.53% | 20.09% | Sr | 20.37% | 19.24% | 21.66%
Se |  | 4 | 3.5 | 3.25 | 4.3333 | 3.25 | 4.4 | 4.2 | 4.3333 | 2.925 | 4.25 | 3.18 | 5 | 4.08 | 3.9333 | 3.4667 |  |  |  | Cd | 56.01% | 42.7% | 43.81% | 77.82% | 41.96% | 22.37% | 56.75% | 65.25% | 51.79% | 62.4% | 69.02% | 46.36% | 73.68% | 23.62% | 33.31% | Cd
Rb |  | 121.4 | 122.6 | 119 | 119.8 | 120.6 | 119.2 | 121.4 | 121.6 | 116.5 | 123.5 | 120.98 | 119.44 | 116.84 | 118.66 | 120.86 |  |  |  | Sb | 371.4% | 229.4% | 214.3% | 237.8% | 276.5% | 251.3% | 242.9% | 288.2% | 329% | 283.4% | 176.2% | 117.6% | 231.4% | 205.2% | 245.3% | Sb
Sr | 242 | 261.8 | 257.6 | 256.4 | 260 | 254.2 | 258.8 | 262.2 | 261.4 | 256.42 | 263.88 | 258.54 | 259.76 | 266.64 | 265.78 | 269.78 |  |  |  | Ba | 83.18% | 78.95% | 61.12% | 90.15% | 78.55% | 115.6% | 68.55% | 58.93% | 97.05% | 57.38% | 79.31% | 107.2% | 79.64% | 46.48% | 61.35% | Ba
Y |  | 32 | 31 | 32.2 | 32.8 | 30.2 | 34 | 32 | 32.4 | 32.18 | 31.14 | 30.98 | 36.18 | 33.5 | 36.34 | 33.82 |  |  |  | Pb | 2.229% | 3.686% | 3.914% | 4.257% | 5.371% | 3.757% | 2.957% | 3.914% | 6.946% | 3.113% | 3.583% | 5.161% | 5.616% | 4.154% | 2.653% | Pb | 4.087% | 3.761% | 4.461%
Zr |  | 290.8 | 317.2 | 305.2 | 306.2 | 300.2 | 294 | 309.2 | 311.8 | 296.58 | 288.26 | 298.58 | 314.4 | 299.62 | 300.86 | 289.68 |  |  |  | U | 32.89% | 232.2% | 66.11% | 165.8% | 57.81% | 32.89% | <LOD | 165.8% | 37.87% | 145.8% | 145.8% | 99.34% | 169.1% | 59.47% | 165.8% | U | 112.6%
Nb |  | 19 | 21 | 19.8 | 18 | 19.2 | 18.6 | 18.6 | 21.6 | 20.44 | 21.66 | 22.2 | 18.82 | 20.98 | 17.76 | 22.2 |  |  |  | average % error -good elements with corrections | 9.647% | 8.994% | 7.775% | 8.644% | 7.392% | 6.624% | 7.254% | 7.202% | 8.087% | 6.755% | 8.57% | 6.299% | 7.404% | 7.566% | 6.589% |  | 7.654% | 8.339% | 7.092%
Mo |  | 11.75 | 10.5 | 9.8 | 6.25 | 8.4 | 10 | 10.2 | 9.25 | 13.06 | 9.44 | 9.4 | 6.96 | 10.525 | 14.4 | 11.02 |  |  |  | average % error - all elements with corrections | 13.82% | 12.71% | 12.34% | 12.7% | 11.53% | 11.08% | 11.05% | 11.4% | 10.62% | 10.14% | 11.06% | 9.078% | 9.878% | 9.825% | 9.07% |  | 11.09% | 11.49% | 9.392%
Ag |  | 24 | 22.5 | 26.6667 | 22.4 | 37.5 | 17.5 | 21.25 | 28.3333 | 21.675 | 22.28 | 30.35 | 25.85 | 23.825 | 15.1 | 31.22 |  |  |  | average % error - elements without corrections | 258.8% | 260.4% | 232.2% | 201% | 168.3% | 203% | 130.2% | 153.9% | 167.4% | 133.7% | 156.1% | 134% | 151.8% | 113.7% | 162.9%
Cd | 54.1 | 84.4 | 77.2 | 77.8 | 96.2 | 76.8 | 66.2 | 84.8 | 89.4 | 82.12 | 87.86 | 91.44 | 79.18 | 93.96 | 66.88 | 72.12 |  |  |  |  | 10% | 10% | 10% | 10% | 10% | 10% | 10% | 10% | 10% | 10% | 10% | 10% | 10% | 10% | 10%
Sn |  | 60.4 | 40.8 | 54.4 | 55.8 | 60.4 | 45.2 | 55.8 | 46.2 | 46.54 | 48.46 | 61.4667 | 56.7667 | 45.375 | 31.28 | 53.58
Sb | 23.8 | 112.2 | 78.4 | 74.8 | 80.4 | 89.6 | 83.6 | 81.6 | 92.4 | 102.1 | 91.26 | 65.725 | 51.78 | 78.88 | 72.64 | 82.18
Ba | 730 | 1048 | 1060.6 | 1168.6 | 1410.2 | 1112.2 | 1310.2 | 1162.6 | 1151.25 | 1106.38 | 1028.12 | 900.05 | 1124.56 | 1282.98 | 808.08 | 932.74
W |  | 30.5 | 21 | 33.8 | 38 | 16 | 36 | 28.6 | 42 | 29.1 | 35.2667 | 25.4333 | 42.5 | 44.55 | 28.9 | 32.025
Au |  | 5.5 | 8 | #DIV/0! | 7.6667 | 10 | #DIV/0! | 5.6667 | 5 | #DIV/0! | #DIV/0! | #DIV/0! | 6.3 | 6.05 | 5.4 | #DIV/0!
Hg |  | 17 | 13.75 | 7.5 | 12.2 | 13.6 | 9.25 | 8 | 11 | 8.3 | 11.65 | 10.22 | 10.7 | 12.2333 | 10.25 | 13.85
Pb | 1400 | 1368.8 | 1348.4 | 1345.2 | 1340.4 | 1324.8 | 1347.4 | 1358.6 | 1345.2 | 1302.76 | 1356.42 | 1349.84 | 1327.74 | 1321.38 | 1341.84 | 1362.86
Bi |  | 28.5 | 18 | 26 | #DIV/0! | 26 | 19.5 | #DIV/0! | 24 | 40.1 | #DIV/0! | #DIV/0! | 27.3 | 26.1 | 28.1 | #DIV/0!
Th |  | 17 | 15.5 | 20 | 10.5 | 12.6667 | 14.5 | 11.75 | 15.3333 | 9.05 | 13 | 16.225 | 14.15 | 15.6 | 12.6 | 15.96
U | 3.01 | 4 | 10 | 5 | 8 | 4.75 | 4 | #DIV/0! | 8 | 4.15 | 7.4 | 7.4 | 6 | 8.1 | 4.8 | 8
LE |  | 729240.2 | 729965.8 | 732139.4 | 730446.2 | 722244.4 | 729122.6 | 725009.6 | 729484 | 742812.06 | 727132.14 | 723277.16 | 725686.66 | 720733.26 | 718122.56 | 719688.14
Mg certified |  | 1.07 | 1.07 | 1.07 | 1.07 | 1.07 | 1.07 | 1.07 | 1.07 | 1.07 | 1.07 | 1.07 | 1.07 | 1.07 | 1.07 | 1.07
1880 m
2610 m